## Slide 1: Lectures on Medical Biophysics
Physical Therapy
Hubbard Hydrotherapy Tank,
Carlos Andreson, Watercolour, 1943

## Slide 2: Lecture outline
Main methods of physical therapy:
Therapy by mechanical treatment
Non-electric thermotherapy – (heating and cooling, hydrotherapy)
Electrotherapy
Ultrasound therapy
Magnetotherapy
Phototherapy

## Slide 3: Therapy by mechanical treatment - examples
Massages – manual or instrumental
Changes in blood circulation, muscular relaxation
Rehabilitative exercises
increase of body strength and mobility, psychical effects, improvement of body posture

## Slide 4: Thermotherapy
The application of heat is (from biophysical point of view) an intervention in the body thermoregulation. Heat can be delivered to the organism (positive thermotherapy), or taken away from the organism (negative thermotherapy).
The body response depends on:
- the way of application - heat conduction, convection or radiation (see electrotherapy and phototherapy)
- the intensity, penetration ability and duration of the heat stimulus. Nonelectric thermotherapy causes mainly changes of body surface temperature (to depths of 2 - 3 cm), with electrotherapy we can heat deeper tissues.
- the size and geometry of the application area in the case of local application: The tissue temperature increases when the heat input from outside exceeds the heat output. Cylindrical body parts are heated faster when the radius is small. Considering only conduction, the resistance to heat flow increases linearly with the thickness of tissue layers. In cylindrically shaped tissues it increases non-linearly.
- the patient’s health (ability of thermoregulation).

## Slide 5: Thermotherapy
The following sources of heat are used in thermotherapy:
Internal (heat produced by the organism itself).
External. Considering the origin and transfer of heat, the thermotherapeutic methods can be divided into five main groups based on:
    - heat conduction
    - heat convection
    - radiation
    - high-frequency electric currents
    - thermal action of ultrasound

## Slide 6: Heat conduction
Mainly packs and compresses. According to the extent of the covered body part, they can be total or partial, according to the temperature hot, indifferent or cold, and wet or dry.
The compresses can be dry (blankets, bottles), peloids (mud) and paraffin. Their temperature ranges from 45 to 55 ºC in dry compresses up to 60 - 77 ºC in paraffin compresses.

## Slide 7: Heat convection – hydrotherapy
hydrotherapy encompasses, besides heat effects, also mechanical action (buoyancy, hydrostatic pressure, impacts of water movement). It acts mainly on the cardiovascular system, vegetative nerves and psychology. Heat helps muscles to relax, reduces pain, accelerates resorption of oedemas. The procedures differ from each other in the way of heat transfer, in the ratio of conduction and convection, and in the degree of homogeneity of heat flux:
cold (less than 18 °C), cool (18 – 24 °C), tepid (24 – 33 °C), warm (33 – 36 °C) or hot (37 – 42 °C).
Or: hypothermic (10 - 34 °C, 5 min.), isothermal (34 - 36°C, 20 - 30 min), hyperthermic (37 - 42°C, short duration).
The effect of the whole-body bath is given mainly by the surface body temperature. After immersion, the body surface is exposed to the actual medium temperature until thermal equilibrium is formed in several millimetres thick water layer, and the effective bath temperature starts to act. Disturbing the layer prevents stabilisation of the effective temperature, thus the patient should not move during the bath.

## Slide 8: Whirling baths, underwater massages, hot and cold water jets
For upper and lower limbs moderately hyperthermic – increasing blood supply and metabolism, skin receptors activated
Alternative application of sharp hot and cold-water jets – a method with outstanding activation effect.
https://www.google.com/url?sa=i&url=https%3A%2F%2Fwww.manes-spa.cz%2Fcs%2Flecba%2Fprocedury&psig=AOvVaw2eGqjcF-ctTDPfDkwC7OUG&ust=1701979704263000&source=images&cd=vfe&opi=89978449&ved=0CBEQjRxqFwoTCNCHqtXO-4IDFQAAAAAdAAAAABA5

## Slide 9: Sauna
Effects of hot (80 - 100°C) air of low relative humidity (10-30%) are utilised, followed by cooling in cold water. Outstanding tonic action. Steam bath: about 45°C, up to 100% humidity. IR sauna

## Slide 10: Cryochamber
Action of dry very cold air (up to -160°C) during some minutes followed by aerobic exercise.

## Slide 11: Cryocauter
Tissue ablation using a stream of cooled media or an instrument tip cooled with e.g. liquid nitrogen, CO2, N2O, special gases-70°C
cervical tumour

## Slide 12: Electrotherapy
Electrotherapeutic methods utilise
Direct electric current (galvanotherapy, iontophoresis)
Low-frequency alternating current or short impulses of direct current (stimulation)
High-frequency alternating current (diathermy)
High frequency electromagnetic radiation

## Slide 13: Conduction of electric currents in tissues
Passage of electric current through human body obeys the Kirchhoff’s laws. Tissue resistance varies. The ions (Na+, K+, Ca2+, Mg2+, Cl- ) are current carriers.
We can distinguish two types of tissue electric conductivity. Cytoplasm and intercellular medium behaves like conductors whose resistance does not depend on frequency. Membrane structures have properties of capacitors, i.e. their impedance Z depends on frequency:

## Slide 14: Resistivity (r, rho) of tissues
R is the electrical resistance of a conductor (measured in Ω)
l is the length of the conductor (measured in m)
A is the cross-sectional area of the conductor (measured in m2)

## Slide 15: Tissue polarisation
The electric charges present in tissues are not always free, they are often bound to macromolecules which are an integral part of cellular structures, and their mobility is limited. The macromolecules behave like electric dipoles – variously oriented – their dipole moments are mutually compensated.
The electric dipoles are oriented according to the direction of the outer electric field when it is present – their polarisation occurs. So an inner electric field of opposite polarity arises, and the intensity of the outer electric field is lowered. This turning of polar molecules gives rise to the so-called displacement current. Permittivity e is a measure of this ability.

## Slide 16: Effects of direct electric current (galvanotherapy, iontophoresis)
Continuous direct current (DC) does not stimulate but can change conditions for that. This effect of DC is called electrotonus and is used in galvanotherapy.
  Around cathode (-) an increase of stimulation of motor nerves occurs = catelectrotonus.
  Around anode (+) a decrease of stimulation of sensitive nerves occurs = anelectrotonus.
  Application in electrotherapy.
Electrokinetic phenomena – movement of ions or solvent in electric field – iontophoresis – ions are transported inside the body.

## Slide 17: Effects of direct electric current
< 1 Khz, pulsed direct current
Pacemaker
Defibrillator (AED)
Electroconvulsive therapy
https://1884403144.rsc.cdn77.org/foto/sarka-rezkova/Zml0LWluLzEyMDB4OTk5OS9maWx0ZXJzOnF1YWxpdHkoODUpOm5vX3Vwc2NhbGUoKS9pbWc/3243382.jpg?v=0&st=oT7E-9ywbgCGxZYWKiYJxCWlwoWGyKTgp_fookcK9QI&ts=1600812000&e=0
https://cdn.hswstatic.com/gif/electroconvulsive-therapy-2-orig.jpg

## Slide 18: Low-frequency AC - electric stimulation
The excitability is a general feature of living systems. In mammals, it is best expressed in nerve and muscle tissue. Electric stimulation - ability of tissue to react on electric stimuli. The direct current has stimulating effects only when suddenly changed.
The stimulation is a threshold phenomenon, it occurs only after a specific threshold intensity has been reached - the rheobase.
The time factor is more important for quantification of stimulation ability: Chronaxie is a time interval necessary for induction of stimulation at the current intensity equal to a two-fold value of rheobase.
Any skeletal muscle has a characteristics chronaxie. Changes of chronaxie help to determine the degree of excitability impairment and also the degree of muscle impairment.
The shortest chronaxie is found in skeletal muscles (< 1 ms), heart muscle (5 ms), the longest one have smooth muscles (50-700 ms). The chronaxie can be read from the so called I/t curve, the dependence of current pulse intensity on its duration.

## Slide 19
The skeletal muscle with normal innervations reacts differently on stimulation by electric impulses with rapid onset (rectangular impulses) and with slow onset (triangular impulses).
In short impulses below about 10 ms, the I/t curve has the same shape.
For longer rectangular impulses the excitability does not change (curve 1) but the excitability for triangular impulses lowers (curve 2).
The muscles with damaged innervations (denervated) are not excitable by very short impulses. However, their excitability for long impulses with slow onset increases (curve 3). So arises area of selective excitability („OSD“), which allows stimulation of denervated muscles without stimulation of healthy muscles.
The I/t curve

## Slide 20: Low-frequency AC - frequency dependence of stimulatory action
In very low frequencies (< 100 Hz), the stimulatory action grows linearly with frequency. In high frequencies, the growth of stimulatory action becomes smaller and changes in decrease. In the range of 500 - 3000 Hz, the threshold value of stimulating current depends on f. The stimulatory action starts to decrease above 3000 Hz and at about 100 kHz disappears fully.
High frequency currents have no stimulatory action because the duration of one period is much shorter than the shortest chronaxie. They have no electrochemical effects either.

## Slide 21
The frequency, length of exposure and direction of passage through the human body are important.The greatest risk of fibrillation is 30-100 Hz Safe current - AC 3.5 mA; DC 10 mA
0.5 to 1 mA - threshold of electric current perception
1 to 8 mA - nerve irritation, blood pressure rise
6 to 15 mA - causes tetanic spasm, person cannot Relax
25 mA - tetanic spasm of the respiratory systém
60 mA - ventricular flutter (fibrillation), transient cardiac arrest
above 80 mA - usually permanent cardiac arrest
above 100 mA - possible burns, defibrillator area
https://jmedicalcasereports.biomedcentral.com/articles/10.1186/s13256-019-2231-4
https://www.akutne.cz/res/algorithm/000105/anot-uraz-elektrickym-proudem_5465c89505bc5.jpg
Negative effects of electric currents
Q = U·I·t

## Slide 22: Electrocauter
Tissue coagulation by thermal effects of an electric current with a frequency of 200 kHz – 3.3 MHz and a power of 750 W (60-120 °C)
Cutting - higher frequency
Coagulation - stop bleeding - lower frequency

## Slide 23: Electrostimulation
The stimulating effects depend on the amplitude, frequency, shape and modulation of pulses, and the kind of tissue!!!!!

## Slide 24: Thermal effects of high frequency (HF) currents
Mechanism of the HF currents action is based on transformation of the absorbed electric energy into heat Q according to Joule’s law:
Q = U·I·t
where U is voltage, t is the time of current I passage. This mechanism of heat production depends on the way of HF currents application.
  Dielectric heating (due to dielectric losses) takes place when applying currents by means of a capacitor field.
  When using induction fields, heat is produced by the so-called eddy currents.

## Slide 25: Utilisation of high-frequency (HF) electric currents
In the case of alternating electric HF currents (>100kHz), the heat effects dominate totally. The heat originates directly in tissues due to dielectric heating, eddy currents or absorption of electromagnetic energy.
For HF therapy, international agreements specified the following frequencies:
  Short-wave diathermy (27.12 MHz, i.e. wavelength of 11.06 m),
  Ultra-short-wave diathermy 433.92 MHz (69 cm),
  Microwave therapy 2 400 or 2 450 MHz (12.4 or 12.25 cm) .
  HF therapy makes possible deep heating.

## Slide 26: Three ways of application of HF currents:
1. The tissue is connected in the electric circuit as a resistor by means of contact electrodes – classical diathermy. It is not used in practice today.
2. Tissue is connected as dielectric placed between insulated electrodes – heating in the capacitor field. The heat produced is proportional to the dielectric loss. Amount of heat arising in subcutaneous fat tissue is lower than in the muscles.
3. Use of eddy currents in magnetic field of a coil – inductive heating. An insulated cable is wound round a limb, or a coil is laid to the body. The skin is less heated, 2 cm thick muscle layer lowers the heating to one half.

## Slide 27: Different ways of HF diathermy
Application of HF currents
(a- condenser field, b- inductive, c- microwaves)

## Slide 28: Short-wave diathermy – heating in capacitor field

## Slide 29: Microwave therapy
Source: magnetron. The oscillations of electromagnetic field are led to an emitter - a dipole with a reflector. 1 cm of muscle is enough to lower the intensity to one half, the relation between heat production in the skin and the muscles is almost equalised. Microwaves put electrically charged particles (ions, dipoles) into oscillatory motion which is transformed into heat by friction.

## Slide 30: Microwave diathermy
(older type)

## Slide 31: Potential risks of microwave and radiofrequency radiation
Mainly thermal effects.
  Microwave sources
  Radars
  Cell phones
  Radio and TV transmitters
  Electric mains
  Trolley lines (wires)
Some studies showing carcinogenic effects of microwaves or low-frequency electromagnetic fields were not verified sufficiently, but it is prudent to reduce exposures.

## Slide 32: Ultrasound therapy
Ultrasound therapy is based on biological effects of ultrasonic oscillations which are not electric. Despite of that, this therapy is sometimes included in the list of electrotherapeutic methods.
An ultrasound (US) therapeutic system consists of two main parts: generator of HF electric current and the application probe, the US source itself, which consists of a piezoelectric transducer.
In therapy, f = 0.8 - 1 MHz is used, sometimes up to 3 MHz, with intensity of US - typically 0.5 - 1 W·cm-2 . Exposure time is 5 - 15 min., in 5 - 10 repetitions. US can be applied continuously or pulsed.
The main therapeutic mechanism is high-frequency massage of tissue. Additional effects are tissue heating (causing hyperaemia) and some physico-chemical effects.
Acoustic coupling between the probe and the skin is secured by a gel (local application) as well as water (underwater application).
Main indications of US therapy: chronic joint, muscle and neural diseases. Limited success is reported in healing wounds after surgery, healing injuries and varicose ulcers.

## Slide 33: Thermal action of ultrasound
In US therapy, thermal dissipation of acoustic energy takes place. Tissue heating depends on physical properties of tissue and its blood supply. The highest heating appears at the interfaces between tissues of very different acoustic impedances.
The thermal action of US cannot be considered without respect to other healing mechanisms (micro-massage etc.)

## Slide 34: Piezoelectric effect
Direct piezoelectric effect
When a piezocrystal is deformed, positive and negative ions in the crystal lattice are displaced and an electric charge is generated on the crystal surface.
The magnitude of the charge is directly proportional to the deformation of the crystal.
The direct effect is used to detect ultrasonic waves.
Indirect piezoelectric effect
By applying an electric voltage to the surface of a piezocrystal, ions in the crystal lattice are displaced and the crystal deforms.
The amount of deformation is directly proportional to the applied voltage.
The indirect effect is used to generate ultrasonic waves.

## Slide 36: Ultrasound - THERAPY

## Slide 37
works on the principle of longitudinal ultrasonic oscillation at a frequency of up to 55 kHz. This generates energy at low temperatures between 50 and 100 °C. It enables cutting and coagulation of soft tissues, blood vessels and sealing of lymphatic pathways in an instant. Bleeding is 4 to 6 times less with the harmonic scalpel than with other surgical methods. Less thermal damage to the tissue.
Ultrasonic Scalpel (harmonic scalpel)

## Slide 38: Cavitation
Ultrasonic cavitation means the disruption of the continuity of the liquid medium by the ultrasonic wave, associated with the formation of gas (cavitation) bubbles.
Resonant cavitation: the bubbles periodically oscillate with the frequency of the ultrasound. The amplitude of the oscillations occurs at a frequency close to the bulk resonance of the bubbles (depending on the size of the bubble).
Collapse cavitation: bubbles collapse rapidly. Their collapse is accompanied by a sharp temperature rise (up to 103 K) and a strong pressure shock (up to 106 Pa). During the collapse, the water vapour in the bubble dissociates to form free radicals.

## Slide 39: Effects of magnetic fields - magnetotherapy
Basic concepts: magnetic fields: static, alternating and pulsed. Homogeneous and non-homogeneous magnetic fields.
Magnetic flux density B depends on the magnetic permeability of the medium m:
m = mr.m0
  Ferromagnetic substances - mr >>1.
  Diamagnetic substances - mr is slightly lower than 1
  Paramagnetic substances - mr is slightly higher than 1.
  (m0 is permeability of vacuum – 4p·10-7 N·A-2)
Body tissues are composed almost only from diamagnetic and paramagnetic substances. Magnetic fields can induce electric voltages and currents in biological medium (due to action of Lorentz force on moving electrical charges, or by action of Faraday force in varying magnetic fields). The induced voltages are, of course, substantially lower than the membrane potentials.

## Slide 40: Magnetomechanical and magnetochemical effects
In a strong mg field, the diamagnetic and paramagnetic molecules orient themselves to minimise their free energy. In non-homogeneous fields with big gradients, a translation movement of ferromagnetic compounds takes place (in living organisms negligible). A strong mg field (over 1 T) would reduce the flow rate of laminar streaming in a tube.
Further, it is necessary to consider indirect action as well, mediated by free radicals arising because of magnetochemical reactions.
We can say that a stable magnetic field of high intensity inhibits metabolic processes, but a varying one stimulates them. These changes are transient.
The interactions of magnetic fields with human tissues are utilised in both diagnostics and therapy. Magnetotherapy is an example of healing procedure. Magnetic stimulation of brain can be used both in diagnostics and therapy.

## Slide 41: Magnets in medicine
Transcranial magnetic stimulation
Magnetotherapy
Biomagnetism quackery of Franz Messmer 200 years ago

## Slide 42: Phototherapy
Ultraviolet (UV), visible (VIS) and infrared (IR) light sources are commonly used in medicine, namely in physical therapy.

## Slide 43: Light radiation
ultraviolet (UV) 1- 380 nm: UV-A 380 - 315 nm
UV-B 315 - 280 nm
UV-C 280 - 190 nm
visible (VIS) 380 - 780 nm
infrared (IR) 0.780 - 1 mm: IR-A 0.78 – 1.4 µm
IR-B 1.4 – 3.0 µm
IR-C 3.0 µm – 1.0 mm
From a practical point of view, the ultraviolet range begins from the wavelength of 190 nm. The spectral range 1 - 190 nm is so called vacuum UV radiation. It is attenuated strongly even by air and hence its biological effects are rare.

## Slide 44: Sources of light
The only natural source is the Sun.
The other sources are artificial and each of them emits only one part of the optical spectrum:
  Hot objects. The wavelength of radiation depends on source temperature, its spectrum is continuous. Light bulbs and various sources of radiant heat.
  Luminescent sources (fluorescent lamps and tubes). They are based on excitation processes in atoms and molecules. Spectrum of these sources can consist of individual spectrum lines.
  Both these sources emit non-coherent radiation.
The only artificial source of intense coherent light is the laser.

## Slide 45: Sources of visible light
http://solarscience.msfc.nasa.gov/predict.shtml

## Slide 46: Molecular mechanisms of biological effects of light
Energy of single atoms depends on its electron configuration. Delivery of energy causes electron jumps to higher energy levels (DEe) – an excited state arises. Absorption spectrum is not continuous. The excitation takes place mainly in the valence shell.
Energetic states DE of a single molecule are, in principle, sums of electron energies DEe corresponding to the electron configuration, vibration energy DEn and rotation energy DEr :
DE = DEe + DEn + Der
All the three kinds of energy are quantised. The action of the radiation depends on photon energy. The lowest energy have photons of IR-C, it corresponds to the changes of rotation energy of molecules. The energy of IR-B and IR-A photons can influence both the vibration and rotation of molecules. The energy of VIS and UV photons can influence rotation, vibration and also electron configuration.

## Slide 47
Interaction of laser radiation with tissue
The effects of laser radiation depend on reflectance, diffusion, absorption and transmission in the biological environment.
The absorbed light laser photon energy is incorporated into the energy processes in the irradiated tissue according to the achieved energy density values and the exposure time.

## Slide 48: Effects of LASER light
During the anti-inflammatory effect, the low-power laser activates the natural repair processes by which the body reacts to the ongoing inflammation. The chemotactic activity of the cells (lymphocytes, monocytes, fibroblasts and polymorphonuclear cells) involved in the anti-inflammatory response of the organism is increased.
The stimulatory effect activates respiratory chain enzymes (flavin dehydrogenase, cytochrome oxidase) in mitochondria.
Enzymes participate in the oxidation reaction resulting in the formation of ATP. Increased concentration of ATP, ADP and AMP, leads to replication of mitochondrial DNA.
The result is accelerated epithelialization of the affected tissue and stimulation of collagen formation.

## Slide 49: Effects of visible light
Photosynthesis biochemistry
Photoreception biophysics of vision
Photosynthesis splits water to liberate O2 and fixes CO2 into sugar
From Wikipedia, the free encyclopedia

## Slide 50: Molecular effects of ultraviolet radiation
Considering compounds of biological importance, the most sensitive are those with conjugated double bonds.
In proteins, the most sensitive amino acids are tyrosine and tryptophan. (abs. maximum around 280 nm).
In NA, the N-bases are sensitive. Their absorbance is higher than the absorbance of proteins, maximum at 240-290 nm.
UV radiation penetrates only into the surface layers of the skin
The skin effect of UV light manifests itself as reddening – erythema – followed by melanin pigmentation  protecting mechanism against further penetration of UV. Synthesis of vitamin D which controls metabolism of Ca and P (its lack causes rickets - rachitis), is an important positive effect of UV light. We cannot also exclude the carcinogenic effect of UV since almost 90% of skin cancer appears on uncovered areas of the skin.

## Slide 51: Sources of ultraviolet radiation
Sun
Mercury discharge tube (used in medicine)
Hydrogen or deuterium discharge tubes (used in research)
Xenon lamp
Electric arc, lightning etc.
Some lasers

## Slide 52: Penetration of UV radiation

## Slide 53: Effects of ultraviolet radiation on living organisms
Sunburns - erythema
Effects on eye: blepharospasm (uncontrollable closure of eye lids) – originates due to damage of cornea by UV radiation.  protection by goggles with UV filter. Lens cataract can arise (Fig.↓)
UV-C with wavelength below 280 nm has outstanding bactericidal action.  sterilisation of labs, special boxes and surgery rooms.

## Slide 54: Sources and effects of infrared light
All the three ranges of IR radiation have thermal effects.
  IR-A is involved in sun light. It passes through glass and is only little absorbed by water.
  IR-B is emitted from various lamps and discharge tubes. It passes through glass but is well absorbed by water.
  IR-C is emitted from heater, hot bodies, humans…. Absorbed by glass and water.
Almost all IR radiation is absorbed by skin. It causes local vasodilatation and thermal erythema which looks like diffuse red patterns and, in contrary to the erythema caused by UV light, its duration is short. Pigmentation is very weak. The irradiation by IR light, however, increases skin sensitivity to the UV light.
A long exposure of eyes to the IR radiation can cause in some professions (glass-blowers, founders, smelters, steelmen etc.) the so called heat cataract (opacity of the lens).

## Slide 55: IR radiation heat transfer
Heat action of the visible and IR light from artificial sources:
Lamp boxes – radiation heat in enclosed space. Skin receptors are stimulated, whole-body heating occurs.
IR lamps: Solux, Sirius - high-power lamps with blue or red filters, radiators of IR light. The radiation is absorbed mainly in body surface. Used mainly in dermatology, ORL and dentistry. Skin receptors are stimulated, suggestive feeling of heat, reflex vasodilatation and muscular relaxation takes place.

## Slide 56: Penetration of IR radiation
Lowering light intensity to 35 % of original value
http://www.depilazione.net/news4.htm

## Slide 57: Thermal Erythema
Thermal erythema – erythema because of excessive use of electric pad -
http://dermatlas.med.jhmi.edu/derm/Display.cfm?ImageName=EAB

## Slide 58: Summary – effects of light

## Slide 59: Author: Vojtěch Mornstein
Content collaboration and language revision: Ivo Hrazdira, Carmel J. Caruana
Last revision September 2024