## Slide 1: FYZIKA VE FIRMĚ Martin Zadražil, 9.11.2011

## Slide 2
Forensic Applications in SEM :
Gun shot residue analysis (GSR)
Bullets and cartridge investigation
After car crash filament and bulb investigation
Tool marks investigation
Analysis of hairs, textiles and papers
Paints, prints and ink analysis
Counterfeit signatures, bank notes
Minerals, soils and metals analysis
and many others...
SEM in Forensic Lab

## Slide 3
Gunshot Residuum analysis.
Identification of the shooter
According GSR particles on hands/clothes
  Typical size, shape
  Typical composition: Pb, Sb, Ba.
Fig. GSR Particles composition in ternary chart
GSR Analysis
Notes: 3

## Slide 4
Value & Excellence in SEMs
Electron Beam Lithography
Beam blanker electro-statically deflects the electron beam
Beam interruption – Beam blanker
0
Beam ON
Blanked beam

## Slide 5
Value & Excellence in SEMs
Technical details - Wide Field OpticsTM

## Slide 6: TESCAN TRACE GSR Hardware
TESCAN TRACE GSR
TESCAN TRACE GSR Hardware Integration

## Slide 7: TESCAN TRACE GSR Hardware
TESCAN TRACE GSR
Powerful 50 MHz internal Pattern Generator
2 x DAC for each stage of deflection coils (X, Y)
16-bit scanning ramp DACs
(65,536 x 65,536 virtual write field)
Variable dwell time
Static and dynamic distortion correction
Digital compensation of field errors
Automatic control of electrostatic beam blanker (10 MHz)
Scanning Generator

## Slide 8
Synthetic YAG crystal
High efficiency – low noise
Fast response
Unlimited lifetime
Suitable for high vacuum
Fig. On-line SE/BSE signal mixing
Dwell/pix | 1Mpix | 100 fields analysis
50 us | 50 s | > 1 hour
100 ns | 0.2 s | < 3 min
Fast detectors save time!!
Fig. Everhart-Thornley SE Detector
SE – Everhardt-Thornley type with YAG
Fast imaging rate (20ns/pix)
First Class Detectors

## Slide 9
BSE Detector
Fig. Phase identification by material contrast
Used for GSR particles location
Tescan R-BSE detector
Fig. Dependecy of BSE intensity on Atomic number
Fig. Thresholding on a callibration sample
Backscattered Electron Detector
Crucial for GSR applications
Detection of particles containing heavy elements

## Slide 10
Fast imaging rates,
but no compromises on Image Quality
Low-end solid-state BSE detector
1us/ pixel
Middle-class scintillation BSE detector
500ns /pxl
Latest generation
Tescan YAG scintillation BSE
100ns /pxl
Testing sample: Gold on Si imaging standard

## Slide 11
Other accessories
EBIC, Absorbed current measurement, IR chamber view camera and many other
EBSD
TEI
EDX/WDX
CL
SEI/BEI
LVSTD
Great Analytical Potential
Great Analytical Potential

## Slide 12
Positioning Accuracy
During GSR analysis the sample is divided into individual fields. Location and analysis of the particles is done filed-by-field using stage movements.
Positioning accuracy is critical parameter for GSR analysis.
Tescan XM Stage:
XY Range : 130 x 130 mm
Min. Step : 300 nm
Relocation accuracy: < 2um guaranteed
Speed: 5 mm/s
Stage movements

## Slide 14
Stage Positioning Accuracy
Image montage from 1048 BSE image fields with two submicron GSR particles found from totaly 48 detected particles.

## Slide 15: Morphological Analysis
TESCAN KOREA Users Meeting 2009, Hanyang University
Roundness: R2/(4*đ*Area)
Circle = 1
Square = 1.3
Triangle = 1.5
Ellipse = 2
Dendrite ~ 15
Thread ~ 25
Extension:
Circle = 0
Square = 0.3
Triangle = 0.5
Ellipse = 2.5
Dendrite ~ 4
Thread ~ 14
Classification by Area
Classification by Roundness
Classification by Extension
(Compactness)
Object size and
shape classification
Over 40 parameters
Object filtering
Statistical output
Notes: 15

## Slide 16: TESCAN ENFSI Results
TESCAN TRACE GSR
TRACE GSR Results
 | Plano GmbH | TESCAN, a.s.
10/15/20 µm | 3 | 3
0,5 µm | 38 | 36
0,7 µm | 30 | 29
1 µm | 33 | 32
1,5 µm | 32 | 30
2 µm | 27 | 27
2,5 µm | 27 | 27
total no. | 190 | 185

## Slide 17: TESCAN ENFSI Results
TESCAN TRACE GSR

## Slide 18
Nanotechnology Workbench
FEG SEM
FIB
GIS
EDS
BSE
EBSD
SE
Nanomanipulators
E-Beam Blanker
Fig. Tescan LYRA Demo as installed on EMC Aachen 2009

## Slide 19
SEM objective lens
FIB column nose
GIS stage with nozzles
What’s Inside?
One Coincidence Point for All Parts

## Slide 20
Ion Beam Imaging
SE generated by ions
Ion imaging resolution < 5 nm
High surface sensitivity with ions
High channeling effect contrast
Ion imaging is destructive!
SE generated by electrons
High Electron Imaging Resolution
Lower surface sensitivity especially for light elements
Nondestructive - reason for “dual beam” systems
Pillar in silicon
diameter 1μm
SE imaging FIB , SEM
Electron Beam Imaging
Why SEM + FIB on One System?

## Slide 21
FIB + GIS Capabilities
Enhanced Milling (Etching) – examples: XeF2 enhances Si, SiO2 and W etching by factor of 5-100 times; H2O of PMMA and diamond 10-20 times
Selective Milling (Etching) – combination of enhanced and reduced milling. Examples: H2O greatly reduces etching of Al, Si and SiO2; XeF2 of Al.
Material deposition – Adsorption of the precursor molecules on the surface, ion beam (or e-beam) induced dissociation of the gas molecules, deposition of the material atoms (e.g. Pt, W, C, SiOx) .
Gas Injection System Option

## Slide 22
Common SEM+FIB Applications
Nano-structuring
IC process inspection and failure analysis
Nano-tomography
Local cross-sectioning, thin film thickness measurement
TEM sample preparation
Biology
Materials science
Mineralogy
Fossils analysis
Forensic applications
Applications

## Slide 23
Forensic Application
Gun Shot Residue Volume Analysis
Simultaneous imaging of the milling process.
Simulteneous Imaging

## Slide 24
Questions ?

## Slide 25
Value & Excellence in SEMs
Applications
Notes: The aim of the NASA’s Genesis sample-return mission was to collect solar wind, material ejected from the outer portion of the sun, and return it to Earth.

## Slide 26
Value & Excellence in SEMs
Applications
The collectors after assembly on Earth before flight
Notes: This was done by allowing the species to implant in the surfaces of ultra-pure collectors. These collectors included single crystal, sapphire, silicon carbide; those materials coated with aluminum, silicon, diamond like carbon, and gold; and isotopically enriched polycrystalline diamond and amorphous carbon. The majority of these materials were distributed on five collector arrays. Three of the materials were housed in an electrostatic concentrator designed to increase the flux of low-mass ions. There was also a two-inch diameter bulk metallic glass collector and a gold foil, polished aluminum, and molybdenum coated platinum foil collector

## Slide 27
Value & Excellence in SEMs
Applications
Genesis Timeline

## Slide 28
Value & Excellence in SEMs
Applications
Notes: The Genesis mission was launched in August 2001. The spacecraft was sent to a special orbit 1 million km from Earth where it was well away from any contamination and exposed ultra-pure collectors to the solar wind. The Genesis mission returned to Earth on September 8, 2004 after a three-year mission.

## Slide 29
Value & Excellence in SEMs
Applications
Notes: The intent was to deploy a drogue chute high over the Utah desert and to catch the capsule in mid-air by helicopter. The capsule would then be opened in a cleanroom constructed for that purpose.

## Slide 30
Value & Excellence in SEMs
Applications
Although the mission was a success with solar wind atoms implanted into the collectors, the return to Earth was ‘non-optimal’ with the parachutes failing to open.
Notes: Unfortunately, both chutes failed to deploy, causing the capsule to fall to the desert floor at a speed of 311 km/h.

## Slide 31
Value & Excellence in SEMs
Applications
Gravity switch
Notes: After studying the problem for a month, the board announced the capsule had crashed because four tiny gravity switches that were supposed to release the parachutes were installed backward.

## Slide 32
Value & Excellence in SEMs
Applications
Monument purchased and designed by Genesis team members to commemorate the return to Earth
Notes: A minute’s of silence shall be observed…..

## Slide 33
Value & Excellence in SEMs
Applications
Hundreds of collector fragments were retrieved

## Slide 34
Value & Excellence in SEMs
Applications
Backside of Collector Array
Gold Foil Collector – post landing

## Slide 35
Value & Excellence in SEMs
Applications
Genesis Timeline

## Slide 36
Value & Excellence in SEMs
Applications
REQUIREMENT> SEM-EDS instrument with the ability to perform fast, automated detection, location and quantification of 10-μm clay-based dirt particles, molybdenum coating and platinum substrate over the entire area of minimum 16 cm x 16 cm square) crumpled Pt-Mo foils. Approximately 8000 cm2 Pt-Mo foils will be imaged in total. Gold Foil Post-landing
< Gold Foil Collector – post landing
Notes: The large samples to be examined are foils (sheets) of Pt with a Mo coating. These foils were used in the Genesis spacecraft to collect solar wind particles; the particles are embedded in the Mo and Au layer. The foils were of course damaged during the crash of the Genesis capsule.
(i) radioactive nuclides in the solar wind collected by the NASA Genesis mission;
(ii) the study of cometary dust collected by the NASA Stardust mission, and
(iii) the study of interstellar dust impacts in foils, also collected by the NASA Stardust mission.

## Slide 37
Value & Excellence in SEMs
Applications
VEGA 3 XMU – X8M5
Notes: Design of the VEGA 3 XMU – X8M5 for the Genesis project

## Slide 38
Value & Excellence in SEMs
Applications
TESCAN's automated particle analysis
Notes: The large samples to be examined are foils of Pt with a Mo coating. Particles are embedded in the Mo layer. The foils were of course damaged during the crash of the Genesis capsule