## Slide 1: Bi9393 Analytická cytometrie Lekce 3
Oddělení cytokinetiky
Biofyzikální ústav AVČR, v.v.i
Královopolská 135
612 65 Brno
e-mail: ksoucek@ibp.cz
tel.: 541 517 166
Karel Souček, Ph.D.
K. Souček Bi9393 Analytická cytometrie

## Slide 2: Principy průtokové cytometrie a sortrování
sorting
zpracování signálu
analýza dat
kompenzace signálu
K. Souček Bi9393 Analytická cytometrie

## Slide 5
+++++
-----

## Slide 6
+++++
-----

## Slide 7
+++++
-----

## Slide 8
+++++
+++++
-----
+++++

## Slide 9
++
+++++
-----
+++++
++
+++++

## Slide 10
+
+++++
-----
+++++
+
+++++

## Slide 11
-----
+++++
-----
+++++
+++++

## Slide 12
-----
+++++
-----
+++++
+++++
-----

## Slide 13
- -
+++++
-----
- - - - -
- -
- - - - -

## Slide 14
-
+++++
-----
-
- - - - -
- - - - -

## Slide 15
-
+++++
-----
- - - - -
- - - - -

## Slide 16
-
+++++
-----
- - - - -
- - - - -

## Slide 17
-
+++++
-----
- - - - -
- - - - -

## Slide 18
http://www.cyto.purdue.edu/cdroms/cyto10a/seminalcontributions/fulwyler.html

## Slide 21
SORTING

## Slide 22
Frequency
Charge
Drop Delay
Amplitude
SORTING

## Slide 23
SORTING
Sheath Pressure
Nozzle Size
Frequency
Amplitude

## Slide 24
SORTING
Drop delay
Interrogation point
Breakoff

## Slide 25
SORTING
Each sort setup includes:
Sheath pressure
Breakoff window values
Side Stream window values
Instrument window > Laser tab values

## Slide 26
SORTING - Streams
Good
Bad

## Slide 27
SORTING – Setup Side Streams

## Slide 28: Drop Delay
interrogation point
drop delay
breakoff
Waste
BD FACS™ Accudrop technology
Accudrop beads
Diode laser
Camera
Optical filter
Notes: Before sorting, you need to make sure that you have an accurate drop delay setting. The Aria has integrated Accudrop technology which provides a simple way to adjust the drop delay.
The components used for Accudrop:
Diode laser: A low-power red laser which illuminates the lower portion of the stream when the sort block door is closed (there’s a safety interlock).
Camera: Provides an image of the side streams visible in this window. (located behind the round window in the sort block)
Optical filter: Used for viewing the fluorescence from the droplets containing Accudrop beads. (point out the button on this image that moves the filter into place)
Accudrop beads: Beads that are excited by the diode laser and emit within the range of the optical filter.
The drop delay is the distance between the laser intercept and the last connected drop, measured in time. The system needs to know when to charge the stream so that the intended droplet is charged before it is detached from the stream.
This diagram shows the diode laser and the lower camera window. When setting the drop delay, the waste aspirator drawer is blocking the collection tubes, so all beads will go to waste. Essentially, Accudrop allows us to view the accuracy of our sort in real-time, without having to do a post-sort analysis.

## Slide 29
Sorting - Sort Masks
Cells are randomized distributed over the stream

## Slide 30
Sorting - Sort Masks
Trailing
Interrogated
Leading

## Slide 31: Mask
A region of the stream monitored for the presence of cells
Determines how drops will be deflected if a sorting conflict occurs
Measured in 1/32 drop increments
Mask = 0
Mask = 8
Mask = 16
Mask = 32
4
4
8
8
16
16
Notes: At the laser intercept particles are defined as wanted or not based on the gates you created in the software. However, while the particle is traveling down the stream another decision is made on whether or not the particle will be sorted based upon the Sort Precision mode.
A sort precision mode is made up of a combination of masks. A mask is a region of the stream that will be considered to make a decision if there’s a sorting conflict. Each mask is measured in 1/32 drop increments.
The next few slides will discuss each individual mask, then we’ll look at how those three masks are combined into precision modes.
Note to instructor: Encourage customers to hold questions about putting the masks together until after you’ve discuss the basic definitions of all three.

## Slide 32: Conflict Resolution
Precision modes include three types of masks
  Yield
  Purity
  Phase

## Slide 33
Sorting - Sort Masks
Sort decisions are determined by sort masks
Target particles in a drop with 1/32-drop resolution

## Slide 34
Sorting - Yield Mask
The yield mask defines how many drops will be sorted Yield mask of 8/32 indicated in blue; target particle shown in green
Yield Mask

## Slide 35
Sorting - Purity Mask
Purity mask of 8/32 in blue, 4/32 in each adjacent drop; target particles in green, non-target particles in red
Purity Mask
Purity Mask

## Slide 37
Flow Cytometry Standard data file format. FCS 3.1
http://www.isac-net.org/images/stories/documents/Standards/fcs3.1_normativespecification_20090813.pdf
Spidlen, J. et al. Cytometry. Part A : the journal of the International Society for Analytical Cytology 77, 97-100, (2010).

## Slide 40: Creation of a Voltage Pulse
Laser
Laser
Laser
Time
Voltage
Time
Voltage
Time
Voltage
1
2
3

## Slide 41: Height, Area, and Width
Time (µs)
Voltage
Pulse area(A)
Pulse Height (H)
Pulse Width (W)
0

## Slide 42: Signal processing
time
analog signal intesity
VOLTAGE
FSC ~ cell size
FL-1 (530/30nm) ~ green fluo.
FL-2 (585/42nm) ~ red fluo.
Analog/digital
conversion
Height
Width
Area ( ∫ )
FL- (H, W, A)
FL-1 (H)
FL-2 (H)
dot plot
0
1000
1000
Zesílení signálu (!)
lin nebo log
K. Souček Bi9393 Analytická cytometrie

## Slide 43
baseline
16, 384
Time
340
1985
7650
12420
Analog to Digital Converter

## Slide 44
Voltage In
PMT Power Supply
Levels 0–1000 Volts
Photon In
Signal
Out
Digital data to memory
Analog to Digital
Conversion
Digitize
the pulse
16,384 levels
Sample
the pulse
10 MHz
Analog to Digital Converter

## Slide 45
Parameters
Area: Sum of all height values
Height: Maximum digitized value X 16
Width: Area/Height X 64K
Data is displayed on 262,144 scale
282
3060
10270
358
4004
9568
14524

## Slide 46: AD převodníky
Počet bitů | # kanálů | rozlišení
8 | 256 | 39.1 mV
10 | 1024 | 9.77 mV
12 | 4096 | 2.44 mV
14 | 16384 | 610 mV
16 | 65536 | 153 mV
18 | 262144 | 38.1 mV
20 | 1048576 | 9.54 mV
22 | 4194304 | 2.38 mV
24 | 16777216 | 596 nV
28 = 256
210 = 1024
.
.
.
Full scale measurement range = 0 to 10 volts
ADC resolution is 12 bits: 212 = 4096 quantization levels
ADC voltage resolution is: (10-0)/4096 = 0.00244 volts = 2.44 mV
K. Souček Bi9393 Analytická cytometrie

## Slide 47: Kolik bitů?
Pokud konvertujeme analogový signál pomocí 8 bitového ADC – máme 256 kanálů (28=256) odpovídajících rozsahu 0-10 V
Rozdíl mezi kanály je 10/256=~40mV
0 50 100 150 200 250
10V
1V
100mV
Channels
upraveno podle J.P.Robinson

## Slide 48: Ideální logaritmický zesilovač
0 50 100 150 200 250
10 V
1 V
100 mV
0 50 100 150 200 250
10 V
1 mV
Channels
Linearní
Log
1 V
100 mV
10 mV
Log amp
upraveno podle J.P.Robinson

## Slide 49: Logaritmické zesílení & dynamický rozsah
upraveno podle J.P.Robinson
lin
log

## Slide 51: Kompenzace fluorescenčního signálu
…později

## Slide 52: Analýza dat
Zobrazení dat
  histogram
  dot plot
  isometric display
  contour plot
  chromatic (color) plots
  3 D projection
Gating

## Slide 53: Způsoby pro zobrazení dat
K. Souček Bi9393 Analytická cytometrie

## Slide 54: Histogram distribuce četnosti
Histogram zobrazuje četnost částic pro jeden parametr
Jednoduchý výstup
Nekorelujeme s dalším parametrem
Problém s identifikací populací
K. Souček Bi9393 Analytická cytometrie

## Slide 55: Dot plot
Zobrazuje korelaci dvou libovolných parametrů
Jednotlivé tečky představují konkrétní změřené buňky (částice)
Hodnoty pro řadu částic mohou ležet ve stejném místě
Nemáme informaci o relativní denzitě částic
Problémy s vykreslením v případě velkých objemů dat
K. Souček Bi9393 Analytická cytometrie

## Slide 56: Density & contour plot
Density plot:
Zobrazuje dva parametry jako frekvenci četnosti
barva a nebo její odstín odpovídá četnosti částic
Contour plot:
spojnice spojuje body (částice) se stejnou hodnotou signálu
V podstatě simulujeme 3D graf – třetí rozměr je frekvence
K. Souček Bi9393 Analytická cytometrie

## Slide 57: Čas jako jeden z parametrů
K. Souček Bi9393 Analytická cytometrie

## Slide 58: 3D zobrazení
2 parametry + četnost
3 parametry společně
K. Souček Bi9393 Analytická cytometrie

## Slide 59: 3 Color Combinations
Negatives
Positives
4+4=8
FITC
PE
APC
4+4+4=12
upraveno podle J.P.Robinson

## Slide 60: 3 Color Combinations
FITC
PE
APC
4+4+4=12
PE-FITC Pattern
FITC
TR-FITC Pattern
FITC
TR-PE Pattern
Phycoerytherin
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
CD4
CD5
CD38
K/L
CD45
CD3
Negative
Negative
CD10
HLA-DR
CD20
CD8
CD7
CD2
CD8
CD8 (dim)
CD2
Negative
CD4
CD5
K
Texas Red
Texas Red
Phycoerytherin
CD3
CD3
IgM
IgG
CD3
CD8
CD3
CD20
IgG
L
CD20
CD20
CD45
upraveno podle J.P.Robinson

## Slide 61: „Gating“
Real-time gating vs. softwarový „gating“
Určení regionů
Strategie „gatingu“
Analýza kvadrantů
Boolean „gating“
zpětný „gating“

## Slide 62: Real-Time vs. Software Gating
Real-time (live) gating:
-omezuje akceptovaná data během měření
Software (analysis) gating:
-vyřazuje určitá data během následné analýzy
upraveno podle J.P.Robinson

## Slide 63: Určení regionů
Objektivní nebo subjektivní?
  -školení/schopnosti/trénink
Možné tvary:
  -obdelník
  -elipsa
  -“free-hand“ (polygon)
  -kvadrant
Statistika
- počet
- podíl (%)
- průměr, medián, S.D., CV, ….

## Slide 64: Region vs. gate
Region
oblast (plocha) v grafu definovaná uživatelem
mnoho regionů v jednou grafu
ohraničujeme pomocí nich populace našeho zájmu
je možné je barevně odlišit
je definován stejně pro všechny vzorky v analýze
Gate
je definován jako jeden a nebo více regionů zkombinovaných pomocí logických operátorů (AND, OR, NOT; Booleova logika)
K. Souček Bi9393 Analytická cytometrie

## Slide 65
Region 1 established
Gated on Region 1
Using Gates
log PE
upraveno podle J.P.Robinson

## Slide 66: Statistika
Aritmerický průměr
Geometrický průměr
Medián
– odhad střední hodnoty
– není ovlivněn extrémními hodnotami
Směrodatná odchylka
Koeficient variance
Modus – nejčastější hodnota
K. Souček Bi9393 Analytická cytometrie

## Slide 67: Statistika pro histogram
K. Souček Bi9393 Analytická cytometrie

## Slide 68: Analýza kvadrantů
(+ +)
( - +)
(+ -)
(- -)
K. Souček Bi9393 Analytická cytometrie

## Slide 69: Logický „Gating“ (Booleova logika)
S překrývajícími se oblastmi máme mnoho možností:
upraveno podle J.P.Robinson

## Slide 70: Boolean Gating
Not Region 1:
upraveno podle J.P.Robinson

## Slide 71: Boolean Gating
Not Region 2:
upraveno podle J.P.Robinson

## Slide 72: Boolean Gating
Region 1 or Region 2:
upraveno podle J.P.Robinson

## Slide 73: Boolean Gating
Region 1 and Region 2:
upraveno podle J.P.Robinson

## Slide 74
Not (Region1 and Region 2):
Boolean Gating
upraveno podle J.P.Robinson

## Slide 75
90 Degree Scatter
0
200
400
600
800
1000
8
15
20
30
40
50
100
200
1000
Lymphocytes
Monocytes
Neutrophils
Side Scatter Projection
Light Scatter Gating
Scale
upraveno podle J.P.Robinson

## Slide 76: Back Gating
Region 4 established
Back-gating using Region 4
log PE
Back gate
upraveno podle J.P.Robinson

## Slide 77: Back Gating
K. Souček Bi9393 Analytická cytometrie

## Slide 78
FITC+
PE+
FITC+
APC+
PE +
APC+
3 Parameter Data Display
upraveno podle J.P.Robinson

## Slide 79: Vícebarevné analýzy generují mnoho dat…
1 2 3 4 5 6 7 8 9 10
3 color
4 color
5 color
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
upraveno podle J.P.Robinson

## Slide 81: Další způsoby vizualizace vícerozměrných dat
t-SNE, viSNE
  t-Distributed Stochastic Neighbor Embedding
  viSNE is a tool for reducing high-parameter data down to two dimensions
  visually identify interesting and rare biological subsets
  allow to gate single cell events across different samples.
SPADE
  Spanning-tree Progression Analysis of Density-normalized Events
  way to automatically identify populations in multidimensional flow cytometry data files
  clusters cells into populations and then projects them into a tree

## Slide 84: The Flow Cytometry: Critical Assessment of Population Identification Methods (FlowCAP)
The goal of FlowCAP is to advance the development of computational methods for the identification of cell populations of interest in flow cytometry data. FlowCAP will provide the means to objectively test these methods, first by comparison to manual analysis by experts using common datasets, and second by prediction of a clinical/biological outcome.

## Slide 85: Způsoby pomocí kterých lze upravit výsledky:
1. Odstranění „doublets“
2. Čas jako parametr pro kontrolu kvality
Příklad - pro DNA analýzu je třeba:
- odstranit „debris“ a shluky
- odstranit „doublets“
- udržovat konstantní průtok
K. Souček Bi9393 Analytická cytometrie

## Slide 86: Co je problém při vícebarevné detekci?
K. Souček Bi9393 Analytická cytometrie

## Slide 87
Emission Spectra–Spectral Overlap

## Slide 88: Kompenzace fluorescenčního signálu při vícebarevné detekci
Proces při kterém dochází k eliminaci všech fluorescenčních signálů kromě signálu z fluorochromu který má být na příslušném detektoru detekován
Nastavení pomocí mixu mikročástic či buněk označených/neoznačených příslušnými fluorochromy.
K. Souček Bi9393 Analytická cytometrie

## Slide 89: Co je problém při vícebarevné detekci?
K. Souček Bi9393 Analytická cytometrie

## Slide 90: Kompenzace fluorescenčního signálu při vícebarevné detekci
K. Souček Bi9393 Analytická cytometrie

## Slide 91
“Bright” = good resolution sensitivity

## Slide 92
Various fluorochromes-stain index

## Slide 93
Choices for 6,- 8,- 10,- and more colors

## Slide 94
Fluorochrome selection considerations

## Slide 95
FITC Spillover
650nm
700nm
PerCP-Cy5.5
695/40
500nm
600nm
FITC
530/30
Relative Intensity
Wavelength (nm)
550nm
PE
585/42
750nm
800nm

## Slide 96
FITC Compensation
650nm
700nm
PerCP-Cy5.5
695/40
500nm
600nm
FITC
530/30
Relative Intensity
Wavelength (nm)
550nm
PE
585/42

## Slide 97
FITC Compensation
650nm
700nm
PerCP-Cy5.5
695/40
500nm
600nm
FITC
530/30
Relative Intensity
Wavelength (nm)
550nm
PE
585/42

## Slide 98
FITC Compensation
Dot plot showing uncompensated FITC data
Dot plot showing compensated FITC data
Biexponential dot plot showing compensated FITC data

## Slide 102: Kompenzace fluorescenčního signálu
K. Souček Bi9393 Analytická cytometrie
FITC positive & negative
PE negative beads
#2

## Slide 103: Kompenzace fluorescenčního signálu
K. Souček Bi9393 Analytická cytometrie
FITC positive & negative
PE negative beads
NONE!

## Slide 104: Kompenzace fluorescenčního signálu

## Slide 105: Nastavení kompenzací
parametr - detektor amp.
FL1 - 544
FL2 - 434
kompenzace
FL1 - 1.1%FL2
FL2 - 17.5%FL1
značené mikročástice – pro běžně konjugované fluorochromy
značené buňky – pro vitální značení

## Slide 106
Effects of Changing PMT Values
FITC Voltage Increased by 5 V
FITC Voltage Decreased by 5 V
Correct Compensation

## Slide 107
Which marker for compensation?
Small errors in compensation of a dim control (A) can result in large compensation errors with bright reagents (B & C).
Use bright markers to setup proper compensation.

## Slide 108
BD Comp Beads
Always positive
Bright staining
Save sample (HIV patients)
Use the same antibody for compensation and the real experiment

## Slide 109
BD Comp Beads

## Slide 110
PBMC were stained as shown in a 3-color experiment. Compensation was properly set for all spillovers
Courtesy Mario Roederer
Fluorescence Minus One

## Slide 111: Tandemové značky

## Slide 112: Tandemové značky - příklad

## Slide 113
Tandems are light sensitive
0 hours
2 hours
22.5 hours
PE
(FL2)
CD8
CD3
PE-Cy5
PE-Cy7
Time Sample Left in Light

## Slide 114: Kompenzace - literatura
Mario Roederer - Compensation in Flow Cytometry
Current Protocols in Cytometry (2002) 1.14.1-1.14.20 John Wiley & Sons, Inc.
M. Loken, D. R. Parks, & L. A. Herzenberg (1977). Two-color immunofluorescence using a fluorescence-activated cell sorter. J. Histochem. Cytochem. 25:899-907. M. Roederer & R. F. Murphy (1986). Cell-by-cell autofluorescence correction for low signal-to-noise systems: application to EGF endocytosis by 3T3 fibroblasts. Cytometry 7:558-565. S. Alberti, D. R. Parks, & L. A. Herzenberg (1987). A single laser method for subtraction of cell autofluorescence in flow cytometry. Cytometry 8:114-119. C. B. Bagwell & E. G. Adams (1993). Fluorescence spectral overlap compensation for any number of flow cytometry parameters. in: Annals of the New York Academy of Sciences, 677:167-184.
K. Souček Bi9393 Analytická cytometrie

## Slide 115
No Data Analysis Technique Can Make Good Data Out of Bad Data!
Shapiro’s 7th Law of Flow Cytometry

## Slide 116: Shrnutí
sorting
zpracování signálu
vizualizace dat a „gating“
kompenzace
Na konci dnešní přednášky byste měli:
Znát základní principu sortování,
popsat způsob zpracování signálu,
rozumět lin / log zesílení signálu,
rozeznat jednotlivé způsoby vizualizace dat,
chápat základní principy „gatingu“,
znát princip kompenzace signálu při vícebarevné detekci.
K. Souček Bi9393 Analytická cytometrie