## Slide 1: Bi9393 Analytická cytometrie
Oddělení cytokinetiky
Biofyzikální ústav AVČR, v.v.i
Královopolská 135
612 65 Brno
e-mail: ksoucek@ibp.cz
tel.: 541 517 166
Karel Souček, Ph.D.
K. Souček Bi9393 Analytická cytometrie

## Slide 2: Fluidika
připomenutí

## Slide 3: Particle Delivery: Hydrodynamic Focusing
Intensity
Count
Narrow particle focus = Narrow distribution
Laser Cross Sectional Area
Sample core is ‘pinched’ by fast flowing sheath
Sample volume ratios of 100 – 1000
Large ratios => low sample inputs
Resolution of particle populations
sheath
sheath
Hydrodynamic core
Conventional Instrumentation: Low Flow Rates (12µL/min)
Notes: To really understand what is different and unique about acoustic focusing, one first needs to understand the current most commonly used method for focusing, called hydrodynamic focusing.
In this method, the sample is injected into a stream of sheath fluid where it is completely surrounded by sheath, the sample core is then centered in the sheath fluid where the laser beam will then interact with the particles. Based on principles relating to laminar flow, the sample core remains separate but coaxial within the sheath fluid. The flow of sheath fluid accelerates the particles and restricts them to the center of the sample core. This process is known as hydrodynamic focusing. Where the sample core is narrow, the particles interact with the laser beam optimally and a narrow distribution results as seen in the histogram above.

## Slide 4: Particle Delivery: Hydrodynamic Focusing
Conventional Instrumentation: High Flow Rate (60µL/min)
Intensity
Count
Broad particle focus = Broad distribution
Increased sample input = increase core size
Particle distributions broadened, CVs increase
Instrument resolution decreased
Historically, low volumetric sample rates used (25 ml/min – 150 ml/min)
sheath
sheath
Hydrodynamic core
Laser Cross Sectional Area

## Slide 5
Attune® Acoustic Focusing Cytometer

## Slide 6: Acoustic Focusing = Better Precision
Acoustic focusing Module
sheath
sheath
sheath
sheath
Narrow particle focus = Narrow distribution
12 µL/min
1000 µL/min
Acoustic focusing of particles occurs prior to mixing with sheath fluid
Notes: Fundamentals of acoustic focusing in flow cytometry:
Picture in the middle is a comparison of a manifold taken from a traditional hydrodynamic focusing cytometer (left) compared to a manifold taken from the Attune Cytometer and represented “in action” by the cartoons directly to the left and right of the photograph. The picture on the far right shows a full acoustic device as well as the manifold at the top.
Here, regardless of the sample volume that is moved through the flow cell, cells are focused by the acoustic device prior to entering the manifold and, ultimately, the flow cell.
So, at low sample volumes (left) and high sample volumes (right), focus is maintained.
This also means that very little sheath is needed. It is only used to allow for varying the amount of sample while maintaining the correct flow rate; and, is used to prevent reagent from staining the inside of the cuvette.

## Slide 7: Principy průtokové cytometrie a sortrování
sorting
zpracování signálu
analýza dat
kompenzace signálu
K. Souček Bi9393 Analytická cytometrie

## Slide 10
+++++
-----

## Slide 11
+++++
-----

## Slide 12
+++++
-----

## Slide 13
+++++
+++++
-----
+++++

## Slide 14
++
+++++
-----
+++++
++
+++++

## Slide 15
+
+++++
-----
+++++
+
+++++

## Slide 16
-----
+++++
-----
+++++
+++++

## Slide 17
-----
+++++
-----
+++++
+++++
-----

## Slide 18
- -
+++++
-----
- - - - -
- -
- - - - -

## Slide 19
-
+++++
-----
-
- - - - -
- - - - -

## Slide 20
-
+++++
-----
- - - - -
- - - - -

## Slide 21
-
+++++
-----
- - - - -
- - - - -

## Slide 22
-
+++++
-----
- - - - -
- - - - -

## Slide 23
http://www.cyto.purdue.edu/cdroms/cyto10a/seminalcontributions/fulwyler.html

## Slide 26
SORTING

## Slide 27
Frequency
Charge
Drop Delay
Amplitude
SORTING

## Slide 28
SORTING
Sheath Pressure
Nozzle Size
Frequency
Amplitude

## Slide 29
SORTING
Drop delay
Interrogation point
Breakoff

## Slide 30
SORTING
Each sort setup includes:
Sheath pressure
Breakoff window values
Side Stream window values
Instrument window > Laser tab values

## Slide 31
SORTING - Streams
Good
Bad

## Slide 32
SORTING – Setup Side Streams

## Slide 33: Drop Delay
interrogation point
drop delay
breakoff
Waste
BD FACS™ Accudrop technology
Accudrop beads
Diode laser
Camera
Optical filter
Notes: Before sorting, you need to make sure that you have an accurate drop delay setting. The Aria has integrated Accudrop technology which provides a simple way to adjust the drop delay.
The components used for Accudrop:
Diode laser: A low-power red laser which illuminates the lower portion of the stream when the sort block door is closed (there’s a safety interlock).
Camera: Provides an image of the side streams visible in this window. (located behind the round window in the sort block)
Optical filter: Used for viewing the fluorescence from the droplets containing Accudrop beads. (point out the button on this image that moves the filter into place)
Accudrop beads: Beads that are excited by the diode laser and emit within the range of the optical filter.
The drop delay is the distance between the laser intercept and the last connected drop, measured in time. The system needs to know when to charge the stream so that the intended droplet is charged before it is detached from the stream.
This diagram shows the diode laser and the lower camera window. When setting the drop delay, the waste aspirator drawer is blocking the collection tubes, so all beads will go to waste. Essentially, Accudrop allows us to view the accuracy of our sort in real-time, without having to do a post-sort analysis.

## Slide 34
Sorting - Sort Masks
Cells are randomized distributed over the stream

## Slide 35
Sorting - Sort Masks
Trailing
Interrogated
Leading

## Slide 36: Mask
A region of the stream monitored for the presence of cells
Determines how drops will be deflected if a sorting conflict occurs
Measured in 1/32 drop increments
Mask = 0
Mask = 8
Mask = 16
Mask = 32
4
4
8
8
16
16
Notes: At the laser intercept particles are defined as wanted or not based on the gates you created in the software. However, while the particle is traveling down the stream another decision is made on whether or not the particle will be sorted based upon the Sort Precision mode.
A sort precision mode is made up of a combination of masks. A mask is a region of the stream that will be considered to make a decision if there’s a sorting conflict. Each mask is measured in 1/32 drop increments.
The next few slides will discuss each individual mask, then we’ll look at how those three masks are combined into precision modes.
Note to instructor: Encourage customers to hold questions about putting the masks together until after you’ve discuss the basic definitions of all three.

## Slide 37: Conflict Resolution
Precision modes include three types of masks
  Yield
  Purity
  Phase

## Slide 38
Sorting - Sort Masks
Sort decisions are determined by sort masks
Target particles in a drop with 1/32-drop resolution

## Slide 39
Sorting - Yield Mask
The yield mask defines how many drops will be sorted Yield mask of 8/32 indicated in blue; target particle shown in green
Yield Mask

## Slide 40
Sorting - Purity Mask
Purity mask of 8/32 in blue, 4/32 in each adjacent drop; target particles in green, non-target particles in red
Purity Mask
Purity Mask

## Slide 42
Flow Cytometry Standard data file format. FCS 3.1
http://www.isac-net.org/images/stories/documents/Standards/fcs3.1_normativespecification_20090813.pdf
Spidlen, J. et al. Cytometry. Part A : the journal of the International Society for Analytical Cytology 77, 97-100, (2010).

## Slide 46: Creation of a Voltage Pulse
Laser
Laser
Laser
Time
Voltage
Time
Voltage
Time
Voltage
1
2
3

## Slide 47: Height, Area, and Width
Time (µs)
Voltage
Pulse area(A)
Pulse Height (H)
Pulse Width (W)
0

## Slide 48: Signal processing
time
analog signal intesity
VOLTAGE
FSC ~ cell size
FL-1 (530/30nm) ~ green fluo.
FL-2 (585/42nm) ~ red fluo.
Analog/digital
conversion
Height
Width
Area ( ∫ )
FL- (H, W, A)
FL-1 (H)
FL-2 (H)
dot plot
0
1000
1000
Zesílení signálu (!)
lin nebo log
K. Souček Bi9393 Analytická cytometrie

## Slide 49
baseline
16, 384
Time
340
1985
7650
12420
Analog to Digital Converter

## Slide 50
Voltage In
PMT Power Supply
Levels 0–1000 Volts
Photon In
Signal
Out
Digital data to memory
Analog to Digital
Conversion
Digitize
the pulse
16,384 levels
Sample
the pulse
10 MHz
Analog to Digital Converter

## Slide 51
Parameters
Area: Sum of all height values
Height: Maximum digitized value X 16
Width: Area/Height X 64K
Data is displayed on 262,144 scale
282
3060
10270
358
4004
9568
14524

## Slide 52: AD převodníky
Počet bitů | # kanálů | rozlišení
8 | 256 | 39.1 mV
10 | 1024 | 9.77 mV
12 | 4096 | 2.44 mV
14 | 16384 | 610 mV
16 | 65536 | 153 mV
18 | 262144 | 38.1 mV
20 | 1048576 | 9.54 mV
22 | 4194304 | 2.38 mV
24 | 16777216 | 596 nV
28 = 256
210 = 1024
.
.
.
Full scale measurement range = 0 to 10 volts
ADC resolution is 12 bits: 212 = 4096 quantization levels
ADC voltage resolution is: (10-0)/4096 = 0.00244 volts = 2.44 mV
K. Souček Bi9393 Analytická cytometrie

## Slide 53: Kolik bitů?
Pokud konvertujeme analogový signál pomocí 8 bitového ADC – máme 256 kanálů (28=256) odpovídajících rozsahu 0-10 V
Rozdíl mezi kanály je 10/256=~40mV
0 50 100 150 200 250
10V
1V
100mV
Channels
upraveno podle J.P.Robinson

## Slide 54: Ideální logaritmický zesilovač
0 50 100 150 200 250
10 V
1 V
100 mV
0 50 100 150 200 250
10 V
1 mV
Channels
Linearní
Log
1 V
100 mV
10 mV
Log amp
upraveno podle J.P.Robinson

## Slide 55: Logaritmické zesílení & dynamický rozsah
upraveno podle J.P.Robinson
lin
log

## Slide 57: Kompenzace fluorescenčního signálu
…později

## Slide 58: Analýza dat
Zobrazení dat
  histogram
  dot plot
  isometric display
  contour plot
  chromatic (color) plots
  3 D projection
Gating

## Slide 59: Způsoby pro zobrazení dat
K. Souček Bi9393 Analytická cytometrie

## Slide 60: Histogram distribuce četnosti
Histogram zobrazuje četnost částic pro jeden parametr
Jednoduchý výstup
Nekorelujeme s dalším parametrem
Problém s identifikací populací
K. Souček Bi9393 Analytická cytometrie

## Slide 61: Dot plot
Zobrazuje korelaci dvou libovolných parametrů
Jednotlivé tečky představují konkrétní změřené buňky (částice)
Hodnoty pro řadu částic mohou ležet ve stejném místě
Nemáme informaci o relativní denzitě částic
Problémy s vykreslením v případě velkých objemů dat
K. Souček Bi9393 Analytická cytometrie

## Slide 62: Density & contour plot
Density plot:
Zobrazuje dva parametry jako frekvenci četnosti
barva a nebo její odstín odpovídá četnosti částic
Contour plot:
spojnice spojuje body (částice) se stejnou hodnotou signálu
V podstatě simulujeme 3D graf – třetí rozměr je frekvence
K. Souček Bi9393 Analytická cytometrie

## Slide 63: Čas jako jeden z parametrů
K. Souček Bi9393 Analytická cytometrie

## Slide 64: 3D zobrazení
2 parametry + četnost
3 parametry společně
K. Souček Bi9393 Analytická cytometrie

## Slide 65: 3 Color Combinations
Negatives
Positives
4+4=8
FITC
PE
APC
4+4+4=12
upraveno podle J.P.Robinson

## Slide 66: 3 Color Combinations
FITC
PE
APC
4+4+4=12
PE-FITC Pattern
FITC
TR-FITC Pattern
FITC
TR-PE Pattern
Phycoerytherin
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
CD4
CD5
CD38
K/L
CD45
CD3
Negative
Negative
CD10
HLA-DR
CD20
CD8
CD7
CD2
CD8
CD8 (dim)
CD2
Negative
CD4
CD5
K
Texas Red
Texas Red
Phycoerytherin
CD3
CD3
IgM
IgG
CD3
CD8
CD3
CD20
IgG
L
CD20
CD20
CD45
upraveno podle J.P.Robinson

## Slide 67: „Gating“
Real-time gating vs. softwarový „gating“
Určení regionů
Strategie „gatingu“
Analýza kvadrantů
Boolean „gating“
zpětný „gating“

## Slide 68: Real-Time vs. Software Gating
Real-time (live) gating:
-omezuje akceptovaná data během měření
Software (analysis) gating:
-vyřazuje určitá data během následné analýzy
upraveno podle J.P.Robinson

## Slide 69: Určení regionů
Objektivní nebo subjektivní?
  -školení/schopnosti/trénink
Možné tvary:
  -obdelník
  -elipsa
  -“free-hand“ (polygon)
  -kvadrant
Statistika
- počet
- podíl (%)
- průměr, medián, S.D., CV, ….

## Slide 70: Region, gate
oblast (plocha) v grafu definovaná uživatelem
mnoho regionů v jednou grafu
ohraničujeme pomocí nich populace našeho zájmu
je možné je barevně odlišit
je definován stejně pro všechny vzorky v analýze
lze je kombinovat pomocí logických operátorů (AND, OR, NOT; Booleova logika)
K. Souček Bi9393 Analytická cytometrie

## Slide 71
Region 1 established
Gated on Region 1
Using Gates
log PE
upraveno podle J.P.Robinson

## Slide 72: Statistika
Aritmerický průměr
Geometrický průměr
Medián
– odhad střední hodnoty
– není ovlivněn extrémními hodnotami
Směrodatná odchylka
Koeficient variance
Modus – nejčastější hodnota
K. Souček Bi9393 Analytická cytometrie

## Slide 73: Statistika pro histogram
K. Souček Bi9393 Analytická cytometrie

## Slide 74: Analýza kvadrantů
(+ +)
( - +)
(+ -)
(- -)
K. Souček Bi9393 Analytická cytometrie

## Slide 75: Logický „Gating“ (Booleova logika)
S překrývajícími se oblastmi máme mnoho možností:
upraveno podle J.P.Robinson

## Slide 76: Boolean Gating
Not Region 1:
upraveno podle J.P.Robinson

## Slide 77: Boolean Gating
Not Region 2:
upraveno podle J.P.Robinson

## Slide 78: Boolean Gating
Region 1 or Region 2:
upraveno podle J.P.Robinson

## Slide 79: Boolean Gating
Region 1 and Region 2:
upraveno podle J.P.Robinson

## Slide 80
Not (Region1 and Region 2):
Boolean Gating
upraveno podle J.P.Robinson

## Slide 81
90 Degree Scatter
0
200
400
600
800
1000
8
15
20
30
40
50
100
200
1000
Lymphocytes
Monocytes
Neutrophils
Side Scatter Projection
Light Scatter Gating
Scale
upraveno podle J.P.Robinson

## Slide 82: Back Gating
Region 4 established
Back-gating using Region 4
log PE
Back gate
upraveno podle J.P.Robinson

## Slide 83: Back Gating
K. Souček Bi9393 Analytická cytometrie

## Slide 84
FITC+
PE+
FITC+
APC+
PE +
APC+
3 Parameter Data Display
upraveno podle J.P.Robinson

## Slide 85
Nástroje pro analýzu dat
Výrobci HW
  Beckman Coulter
    Kaluza
  Becton Dickinson
    FACSDiva
    FACSSuite
    FlowJo
  BioRad
  Sony
  Milteney
  …
Univerzální platformy
  Komerční
    FlowJo
    FCS Express
    …
  Freeware
    Flowing Software
    Cyflogic
    BioConductor - Flowcore

## Slide 86: Vícebarevné analýzy generují mnoho dat…
1 2 3 4 5 6 7 8 9 10
3 color
4 color
5 color
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
upraveno podle J.P.Robinson

## Slide 88: Další způsoby vizualizace vícerozměrných dat
t-SNE, viSNE
  t-Distributed Stochastic Neighbor Embedding
  viSNE is a tool for reducing high-parameter data down to two dimensions
  visually identify interesting and rare biological subsets
  allow to gate single cell events across different samples.
SPADE
  Spanning-tree Progression Analysis of Density-normalized Events
  way to automatically identify populations in multidimensional flow cytometry data files
  clusters cells into populations and then projects them into a tree

## Slide 91: The Flow Cytometry: Critical Assessment of Population Identification Methods (FlowCAP)
The goal of FlowCAP is to advance the development of computational methods for the identification of cell populations of interest in flow cytometry data. FlowCAP will provide the means to objectively test these methods, first by comparison to manual analysis by experts using common datasets, and second by prediction of a clinical/biological outcome.

## Slide 92: Způsoby pomocí kterých lze upravit výsledky:
1. Odstranění „doublets“
2. Čas jako parametr pro kontrolu kvality
Příklad - pro DNA analýzu je třeba:
- odstranit „debris“ a shluky
- odstranit „doublets“
- udržovat konstantní průtok
K. Souček Bi9393 Analytická cytometrie

## Slide 93
https://www.flowjo.com/learn/flowjo-university/flowjo/tutorial/33

## Slide 94: Co je problém při vícebarevné detekci?
K. Souček Bi9393 Analytická cytometrie

## Slide 95
Emission Spectra–Spectral Overlap

## Slide 96: Kompenzace fluorescenčního signálu při vícebarevné detekci
Proces při kterém dochází k eliminaci všech fluorescenčních signálů kromě signálu z fluorochromu který má být na příslušném detektoru detekován
Nastavení pomocí mixu mikročástic či buněk označených/neoznačených příslušnými fluorochromy.
K. Souček Bi9393 Analytická cytometrie

## Slide 97: Co je problém při vícebarevné detekci?
K. Souček Bi9393 Analytická cytometrie

## Slide 98: Kompenzace fluorescenčního signálu při vícebarevné detekci
K. Souček Bi9393 Analytická cytometrie

## Slide 99
“Bright” = good resolution sensitivity

## Slide 100
Various fluorochromes-stain index

## Slide 101
Choices for 6,- 8,- 10,- and more colors

## Slide 102
Fluorochrome selection considerations

## Slide 103
FITC Spillover
650nm
700nm
PerCP-Cy5.5
695/40
500nm
600nm
FITC
530/30
Relative Intensity
Wavelength (nm)
550nm
PE
585/42
750nm
800nm

## Slide 104
FITC Compensation
650nm
700nm
PerCP-Cy5.5
695/40
500nm
600nm
FITC
530/30
Relative Intensity
Wavelength (nm)
550nm
PE
585/42

## Slide 105
FITC Compensation
650nm
700nm
PerCP-Cy5.5
695/40
500nm
600nm
FITC
530/30
Relative Intensity
Wavelength (nm)
550nm
PE
585/42

## Slide 106
FITC Compensation
Dot plot showing uncompensated FITC data
Dot plot showing compensated FITC data
Biexponential dot plot showing compensated FITC data

## Slide 110: Kompenzace fluorescenčního signálu
K. Souček Bi9393 Analytická cytometrie
FITC positive & negative
PE negative beads
#2

## Slide 111: Kompenzace fluorescenčního signálu
K. Souček Bi9393 Analytická cytometrie
FITC positive & negative
PE negative beads
NONE!

## Slide 112: Kompenzace fluorescenčního signálu

## Slide 113: Nastavení kompenzací
parametr - detektor amp.
FL1 - 544
FL2 - 434
kompenzace
FL1 - 1.1%FL2
FL2 - 17.5%FL1
značené mikročástice – pro běžně konjugované fluorochromy
značené buňky – pro vitální značení

## Slide 114
Effects of Changing PMT Values
FITC Voltage Increased by 5 V
FITC Voltage Decreased by 5 V
Correct Compensation

## Slide 115
Which marker for compensation?
Small errors in compensation of a dim control (A) can result in large compensation errors with bright reagents (B & C).
Use bright markers to setup proper compensation.

## Slide 116
BD Comp Beads
Always positive
Bright staining
Save sample (HIV patients)
Use the same antibody for compensation and the real experiment

## Slide 117
BD Comp Beads

## Slide 118
PBMC were stained as shown in a 3-color experiment. Compensation was properly set for all spillovers
Courtesy Mario Roederer
Fluorescence Minus One

## Slide 119: Tandemové značky

## Slide 120: Tandemové značky - příklad

## Slide 121
Tandems are light sensitive
0 hours
2 hours
22.5 hours
PE
(FL2)
CD8
CD3
PE-Cy5
PE-Cy7
Time Sample Left in Light

## Slide 122: Kompenzace - literatura
Mario Roederer - Compensation in Flow Cytometry
Current Protocols in Cytometry (2002) 1.14.1-1.14.20 John Wiley & Sons, Inc.
M. Loken, D. R. Parks, & L. A. Herzenberg (1977). Two-color immunofluorescence using a fluorescence-activated cell sorter. J. Histochem. Cytochem. 25:899-907. M. Roederer & R. F. Murphy (1986). Cell-by-cell autofluorescence correction for low signal-to-noise systems: application to EGF endocytosis by 3T3 fibroblasts. Cytometry 7:558-565. S. Alberti, D. R. Parks, & L. A. Herzenberg (1987). A single laser method for subtraction of cell autofluorescence in flow cytometry. Cytometry 8:114-119. C. B. Bagwell & E. G. Adams (1993). Fluorescence spectral overlap compensation for any number of flow cytometry parameters. in: Annals of the New York Academy of Sciences, 677:167-184.
K. Souček Bi9393 Analytická cytometrie

## Slide 123
No Data Analysis Technique Can Make Good Data Out of Bad Data!
Shapiro’s 7th Law of Flow Cytometry

## Slide 124: Shrnutí
sorting
zpracování signálu
vizualizace dat a „gating“
kompenzace
Na konci dnešní přednášky byste měli:
Znát základní principu sortování,
popsat způsob zpracování signálu,
rozumět lin / log zesílení signálu,
rozeznat jednotlivé způsoby vizualizace dat,
chápat základní principy „gatingu“,
znát princip kompenzace signálu při vícebarevné detekci.
K. Souček Bi9393 Analytická cytometrie