## Slide 1: F9180: Diagnostické metody 2
Time-Correlated Single Photon Counting Methods
Tomáš Hoder
Doporučená literatura:
[1] D.V.O’Connor and D.Phillips - Time-correlated Single Photon Counting, 1984
[2] W.Demtroeder - Atoms, Molecules and Photons, 2006
[3] W.R.Ware – Techniques of pulse fluorometry Time-Resolved Fluorescence
Spectroscopy in Biochemistry and Biology (NATO ASI Series A: Life Sciences)
vol 69, ed R.B.Cundall and R.E.Dale (New York: Plenum), 1983
[4] K.V.Kozlov et al. 2001 Spatio-temporally resolved spectroscopic diagnostics
of the barrier discharge in air at atmospheric pressure, J.Phys.D:Appl.Phys. 34 3164
[5] W.Becker 2007 Advanced Time-correlated Single Photon Counting Techniques
[6] W.Becker 2014 TCSPC Handbook

## Slide 2: Overview
TC-SPC technique/idea
Time scales, where we can use it, where not
Light emission, fluorescence, quenching
Sensitivity, StNR, time resolution, PMT transit
Use in plasma-physics and synchronization
Selected examples for gas discharge quantitative spectroscopy

## Slide 3: TC-SPC technique
SYNC-Signal:
represents shape of the full light pulse
giving a time reference
START
STOP
Delay
single photon
accumulation
MAIN-Signal:
random single photons
spatially resolved
spectrally resolved
micro-discharge
first counted photon
time
relative
time information
triggering possible on stochastically appearing events in time as the time reference is set on the discharge itself >>> cross-correlation spectroscopy

## Slide 4: TC-SPC basics
Becker 2006 Advanced TCSPC techniques
Instead of having problems with slow
analogue PMT signal …
TC-SPC obtains light intensity by
counting pulses as a digital units in
subsequent time channels:
free of gain noise of PMT
free of electronic noise of accidental signals to PMT
high signal-to-noise ratio (to PMT background counting rate)
higher time-resolution (Transit Time Spread << Single Photoelectron Response/Transit time)

## Slide 5: Basics

## Slide 6: TC-SPC solves:
Problems of triggering/synchro – recording of irregular emission events
Problems of time resolution – no limitation by transit times or SER (single electron response) of detectors
Problems of sensitivity – statistical principles behind the accumulative recordings

## Slide 7: Convolution -> cross-correlation
wiki
(δ-fce)

## Slide 8: Relevant integral transformations:
wiki

## Slide 9: Different time scales and appropriate light detection techniques
Gated cameras:
Time resolution from seconds to usually 2ns (new models down to 200 or even 50ps)
Almost impossible to synchronize to time-irregular emission events (random shot is the time consuming solution)
Sensitivity of the recording is given by the StNR of given device. Usually the noise increases linearly with number of accumulation cycles. Weak signals are not easy to record.
Notes: EEDF relaxation ps … radiative lifetimes 10ns … effective quenching at 1atm ns … relaxation of chemical kinetics ns … streamer structure 10ps …

## Slide 10: Different time scales and appropriate light detection techniques
Start-stop TCSPC:
Time resolution from hundreds of nanoseconds to 10ps
Possible to synchronize to time-irregular emission events with the same high resolution in time
Sensitivity of the recording is given by the StNR of given synchronization arrangement. Poisson statistics is the limiting mechanism.
Limited for high-frequency repetition emission events. Limited by the speed of electronics of the counter dealing with high-frequency input.
Time-to-amplitude converter
… first time 1961 by Koechlin (Thesis, Uni Paris)

## Slide 11: Different time scales and appropriate light detection techniques
Reversed start-stop TCSPC:
Time resolution from hundreds of nanoseconds to 10ps
Possible to synchronize to time-irregular emission events
Sensitivity of the recording is given by the StNR of given synchronization arrangement. Poisson statistics is the limiting mechanism.
No limitation for high-frequency repetition emission events. Input processing only for the “main” signal.
SYNC-Signal:
represents shape of the full light pulse
giving a time reference
START
STOP
Delay
MAIN-Signal:
random single photons
spatially resolved
spectrally resolved
micro-discharge

## Slide 12: Different time scales and appropriate light detection techniques
Streak cameras:
Time resolution down to units of picoseconds (some new models down to hundreds of femtoseconds)
Sensitivity of the recording comparable to the TC-SPC

## Slide 13: Different time scales and appropriate light detection techniques
Pump-and-probe technique:
Time resolution down to femtoseconds using femtosecond laser pulses
Possible to synchronize due to the synchronous generation of the fluorescence by the pumping laser pulse.

## Slide 14: Light emission, fluorescence
Notes: The main reason and use of TCSPC is in fluorescence lifetime imaging microscopy >>> proto ted lifetime, je to potřeba i ve fyzice plazmatu

## Slide 15: Light emission, fluorescence

## Slide 16: Light emission, fluorescence

## Slide 17: Radiative lifetime measurement
(j = n)
(k = n)

## Slide 18: Effective lifetime and quenching

## Slide 19: Effective lifetime and quenching
Dilecce et al. 2010 J.Phys.D

## Slide 20: TC-SPC technique
SYNC-Signal:
represents shape of the full light pulse
giving a time reference
START
STOP
Delay
single photon
accumulation
MAIN-Signal:
random single photons
spatially resolved
spectrally resolved
micro-discharge
first counted photon
time
relative
time information
triggering possible on stochastically appearing events in time as the time reference is set on the discharge itself >>> cross-correlation spectroscopy

## Slide 21: Convolution >>> cross-correlation
wiki
(δ-fce)

## Slide 22: TCSPC statistics scheme
NE
zi
wi
NA
TAC
Ni=NA
i-th channel
Ni
ND
NA number of counts (anode pulses) in i-th interval
TAC detekuje jen první foton
a tedy NA Ni !
ND – total number of anode pulses for all channels,
i.e.

## Slide 23: TC-SPC statistics basics 1
Photoelectrons generated by impinged
photons on the cathode with given quantum efficiency.
The probability of emission of l
photoelectrons in the i-th interval
is given by the Poisson distribution.
The Taylor series of the exponential function is (1-wi+wi2/2+…),
we take the first two.

## Slide 24: TCSPC statistics scheme
NE
zi
wi
NA
TAC
Ni=NA
i-th channel
Ni
ND
NA number of counts in i-th interval

## Slide 25: TC-SPC statistics basics 2
After developing to Taylor series,
as shown before:
And it follows:
After a large number of excitation pulses NE,
the number of anode pulses NA in the i-th interval.
Therefore the number of anode pulses NA is
proportional to the intensity of the fluorescence
at time ti.

## Slide 26: TCSPC statistics scheme
NE
zi
wi
NA
TAC
Ni=NA
i-th channel
Ni
ND
NA number of counts in i-th interval

## Slide 27: TC-SPC statistics basics 3
Relation of NA to number of
Counts in the i-th channel Ni:
Because the TAC detects only the first photon
in given time interval for a given excitation cycle,
NA is not the number of counts in the i-th channel Ni.
The true relation is given left.
Consequently the count in channel i is a measure
of the fluorescence intensity at time ti.

## Slide 28: TC-SPC statistics basics 4
PILE-UP effect

## Slide 29: Other issues to be aware of
Color effect (consequence of photoeffect)
Afterpulsing (consequence of PMT setup)
Ultra-short reflections
…

## Slide 30: Sensitivity and precision
The effect of PMT noise is greatly reduced by the mode of TAC operation
>>> enhanced Signal-to-Noise ratio (up to 100x noise reduction)
Noise due to the dark counts on PMT (cooling, background subtraction …)
Noise due to the counting error, number of counts in each channel I(ti) follows
a Poisson distribution with a standard deviation σi given by σi=(I(ti))1/2
It follows that to have 5% precision in the number of Ni counts in i-th channel,
where the the curve decayed to 1% of its maximum value, one has: 0.05=1/σi =1/(Ni)1/2
and Ni is 400, that means one has to measure 40000 counts in maximum
Signal-to-Noise ratio is given as well by the Poisson distribution and is equal
to the standard deviation:
Dynamic range (ratio between the largest and smallest value of measured quantity):
for ICCD typically 1000:1, streak 10000:1, for TCSPC usually 100000:1 and more

## Slide 31: Comparison
Hamamatsu News 2009

## Slide 32: PMT and MCP structure
Gain of up to 108

## Slide 33: PMT resolution, transit time
Becker 2006 Advanced TCSPC techniques
Leading edge discrimination
CFD
pulse-height-induced timing jitter avoided

## Slide 34: PMT resolution, transit time
Becker 2006 Advanced TCSPC techniques

## Slide 35: PMT transit time spread

## Slide 36: PMT transit
Lakowicz 2006 Principles of fluorescence spectroscopy

## Slide 37: TC-SPC review
Becker 2006 Advanced TCSPC techniques
light intensity > 0.01 to 0.1 photons per signal period  Pile-Up Problem
Relatively slow recording speed and long data acquisition times  high repetition rates and low dead time (approx. 100 ns; i.e. 107 photons/s)

## Slide 38: Use in plasma-physics and signal synchronization
Short discharges with high repetition:
rf discharges, barrier discharges, Trichel
pulsing corona, self-pulsing
sparks
Synchronization via light pulse, current pulse, laser excitation or TTL of applied voltage waveform

## Slide 39: Kinetic scheme dependent example
Streamer discharges generated in atmospheric pressure air
Spectra is dominated by the second positive system of molecular nitrogen
Relatively weak bands of first negative system are present as well

## Slide 40: Kinetic scheme dependent example
Streamer discharges generated in atmospheric pressure air
Spectra is dominated by the second positive system of molecular nitrogen
Relatively weak bands of first negative system are present as well

## Slide 41: Kinetic scheme dependent example
Streamer discharges generated in atmospheric pressure air

## Slide 42: Trichel pulse corona
Breakdown in negative corona Trichel pulse

## Slide 43: Setup for corona and calibration

## Slide 44: Calibration
Streamer discharges generated in atmospheric pressure air

## Slide 45: Trichel pulse electric field
SPS
FNS

## Slide 46: Limits for TCSPC on streamers

## Slide 47: Quantum mechanics based example
Obradovic 2008 APL

## Slide 48: Quantum mechanics based example
Kuraica 1997 APL

## Slide 49: RF discharge in helium at 1atm
Navratil 2015 TCSPC results

## Slide 50: Quantum mechanics based example
Navratil 2015 TCSPC results

## Slide 51: Summary
The principles and technical realization of the TCSPC technique were introduced
The measured fluorescence from the atomic/molecular transitions was followed back to its origin
Selected examples of quantitative high-resolution spectroscopy were presented