|  | www.qiagen.com |
| --- | --- |

# Quantitation Report

#### Experiment Information

| Run Name | Run 2021-12-15-sk3 |
| --- | --- |
| Run Start | 15.12.2021 10:53:27 |
| Run Finish | 15.12.2021 12:30:12 |
| Operator | MP |
| Notes |  |
| Run On Software Version | Rotor-Gene Q Software 2.0.2.4 |
| Run Signature | The Run Signature is valid. |
| Gain Green | 10, |
| Machine Serial No. | 0911113 |

#### Quantitation Information

| Threshold | 0,030 |
| --- | --- |
| Left Threshold | 5,000 |
| Standard Curve Imported | No |
| Standard Curve (1) | conc= 10^(-0,360*CT + 10,759) |
| Standard Curve (2) | CT = -2,779*log(conc) + 29,898 |
| Reaction efficiency (*) | (* = 10^(-1/m) - 1) 1,29001 |
| M | -2,77904 |
| B | 29,89843 |
| R Value | 0,99501 |
| R^2 Value | 0,99004 |
| Start normalising from cycle | 1 |
| Noise Slope Correction | Yes |
| No Template Control Threshold | % 0 |
| Reaction Efficiency Threshold | Disabled |
| Normalisation Method | Dynamic Tube Normalisation |
| Digital Filter | Light |
| Sample Page | Page 2 |
| Imported Analysis Settings |  |

#### Profile

| Cycle | Cycle Point |
| --- | --- |
| Hold 1 | Hold @ 95°C, 10min 0s |
| Cycling (35 repeats) | Step 1: Hold @ 95°C, 15s |
|  | Step 2: Hold @ 60°C, 30s, acquiring to Cycling A([Green][1][1]) |
|  | Step 3: Hold @ 72°C, 20s |
| Hold 2 | Hold @ 72°C, 1min 0s |
| Melt | Ramp from 55°C to 90°C |
|  | Hold for 90s on the 1st step |
|  | Hold for 5s on next steps,Melt A([Green][1][1]) |

#### Raw Data For Cycling A.Green

#### Quantitation data for Cycling A.Green

#### Standard Curve

| No. | Color | Name | Type | Ct | Ct Comment | Given Conc (IU/ml) | Calc Conc (IU/ml) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 41 |  | T6 3 1X | Standard | 18,71 |  | 10 000 | 10 593 |
| 42 |  | T6 3 1X | Standard | 19,37 |  | 10 000 | 6 147 |
| 43 |  | T6 3 10x | Standard | 20,99 |  | 1 000 | 1 603 |
| 44 |  | T6 3 10x | Standard | 21,05 |  | 1 000 | 1 525 |
| 45 |  | T6 3 100x | Standard | 24,83 |  | 100 | 67 |
| 47 |  | T6 3 1000x | Standard | 27,20 |  | 10 | 9 |
| 48 |  | T6 3 1000x | Standard | 27,22 |  | 10 | 9 |
| 49 |  | T6 3 10000x | Standard | 29,79 |  | 1 | 1 |
| 51 |  | 3 NK | Unknown | 31,76 |  |  |  |
| 52 |  | 3 NK | Unknown |  |  |  |  |
| 53 |  | D6 3 ACO4 | Unknown |  |  |  |  |
| 54 |  | D6 3 ACO4 | Unknown |  |  |  |  |
| 55 |  | D6 3 ACO4 | Unknown | 28,89 |  |  | 2 |
| 56 |  | 3 NK | Unknown | 34,85 |  |  |  |
| 58 |  | B6 3 ACO4 | Unknown | 15,87 |  |  | 111 541 |
| 59 |  | B6 3 ACO4 | Unknown | 15,17 |  |  | 199 580 |
| 60 |  | 3 NK | Unknown | 29,68 |  |  | 1 |

-

Warning: The following samples were not analysed :
1T6 1 1X- 2T6 1 1X- 3T6 1 10x- 4T6 1 10x- 5T6 1 100x- 6T6 1 100x- 7T6 1 1000x- 8T6 1 1000x- 9T6 1 10000x- 10T6 1 10000x- 111 NK- 121 NK- 13D6 1 ACO3- 14D6 1 ACO3- 15D6 1 ACO3- 16B6 1 ACO3- 17B6 1 ACO3- 18B6 1 ACO3- 191 NK- 201 NK- 21T6 2 10X- 22T6 2 10X- 23T6 2 1x- 24T6 2 1x- 25T6 2 100x- 26T6 2 100x- 27T6 2 1000x- 28T6 2 1000x- 29T6 2 10000x- 30T6 2 10000x- 312 NK- 322 NK- 33D6 2 ACO2- 34D6 2 ACO2- 35D6 2 ACO2- 362 NK- 37B6 2 ACO2- 38B6 2 ACO2- 39B6 2 ACO2- 402 NK- 46T6 3 100x- 50T6 3 10000x- 57B6 3 ACO4

Legend:
NEG (NTC) - Sample cancelled due to NTC Threshold.
NEG (R. Eff) - Sample cancelled as efficiency less than reaction efficiency threshold.

This report was generated by Rotor-Gene Q Series Software 2.3.1 (Build 49) Copyright 2013 QIAGEN GmbH. All Rights Reserved.