|  | www.qiagen.com |
| --- | --- |

# Quantitation Report

#### Experiment Information

| Run Name | Run 2021-12-15-sk4 |
| --- | --- |
| Run Start | 15.12.2021 13:53:27 |
| Run Finish | 15.12.2021 15:40:16 |
| Operator | MP |
| Notes |  |
| Run On Software Version | Rotor-Gene Q Software 2.0.2.4 |
| Run Signature | The Run Signature is valid. |
| Gain Green | 10, |
| Machine Serial No. | 0911113 |

#### Quantitation Information

| Threshold | 0,030 |
| --- | --- |
| Left Threshold | 5,000 |
| Standard Curve Imported | No |
| Standard Curve (1) | conc= 10^(-0,462*CT + 13,104) |
| Standard Curve (2) | CT = -2,165*log(conc) + 28,364 |
| Reaction efficiency (*) | (* = 10^(-1/m) - 1) 1,89724 |
| M | -2,16458 |
| B | 28,36425 |
| R Value | 0,94975 |
| R^2 Value | 0,90203 |
| Start normalising from cycle | 1 |
| Noise Slope Correction | Yes |
| No Template Control Threshold | % 0 |
| Reaction Efficiency Threshold | Disabled |
| Normalisation Method | Dynamic Tube Normalisation |
| Digital Filter | Light |
| Sample Page | Page 2 |
| Imported Analysis Settings |  |

#### Profile

| Cycle | Cycle Point |
| --- | --- |
| Hold 1 | Hold @ 95°C, 10min 0s |
| Cycling (40 repeats) | Step 1: Hold @ 95°C, 15s |
|  | Step 2: Hold @ 60°C, 30s, acquiring to Cycling A([Green][1][1]) |
|  | Step 3: Hold @ 72°C, 20s |
| Hold 2 | Hold @ 72°C, 1min 0s |
| Melt | Ramp from 55°C to 90°C |
|  | Hold for 90s on the 1st step |
|  | Hold for 5s on next steps,Melt A([Green][1][1]) |

#### Raw Data For Cycling A.Green

#### Quantitation data for Cycling A.Green

#### Standard Curve

| No. | Color | Name | Type | Ct | Ct Comment | Given Conc (IU/ml) | Calc Conc (IU/ml) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  | T6 1X | Standard | 19,08 |  | 10 000 | 19 419 |
| 2 |  | T6 1X | Standard | 18,96 |  | 10 000 | 22 162 |
| 3 |  | T6 10x | Standard | 21,86 |  | 1 000 | 1 014 |
| 4 |  | T6 10x | Standard | 22,61 |  | 1 000 | 458 |
| 5 |  | T6 100x | Standard | 24,96 |  | 100 | 37 |
| 6 |  | T6 100x | Standard | 25,42 |  | 100 | 23 |
| 7 |  | T6 1000x | Standard | 24,55 |  | 10 | 58 |
| 8 |  | T6 1000x | Standard | 26,55 |  | 10 | 7 |
| 9 |  | T6 10000x | Standard | 29,41 |  | 1 |  |
| 10 |  | T6 10000x | Standard | 26,96 |  | 1 | 4 |
| 11 |  | NK | Unknown | 26,99 |  |  | 4 |
| 12 |  | NK | Unknown | 24,67 |  |  | 51 |
| 13 |  | D6 ACO1 | Unknown | 15,83 |  |  | 614 583 |
| 14 |  | D6 ACO1 | Unknown | 15,39 |  |  | 988 758 |
| 15 |  | D6 ACO1 | Unknown | 15,43 |  |  | 949 130 |
| 16 |  | B6 ACO1 | Unknown | 14,86 |  |  | 1 734 706 |
| 17 |  | B6 ACO1 | Unknown | 14,91 |  |  | 1 637 042 |
| 18 |  | B6 ACO1 | Unknown | 14,82 |  |  | 1 810 215 |
| 19 |  | NK | Unknown | 28,58 |  |  | 1 |
| 20 |  | NK | Unknown | 30,22 |  |  |  |

-

Warning: The following samples were not analysed :
21T6 1X- 22T6 1X- 23T6 10x- 24T6 10x- 25T6 100x- 26T6 100x- 27T6 1000x- 28T6 1000x- 29T6 10000x- 30T6 10000x- 31NK- 32NK- 33D6 ACO5- 34D6 ACO5- 35D6 ACO5- 36B6 ACO5- 37B6 ACO5- 38B6 ACO5- 39NK- 40NK- 41T6 1X- 42T6 1X- 43T6 10x- 44T6 10x- 45T6 100x- 46T6 100x- 47T6 1000x- 48T6 1000x- 49T6 10000x- 50T6 10000x- 51NK- 52NK- 53D6 UBC10- 54D6 UBC10- 55D6 UBC10- 56B6 UBC10- 57B6 UBC10- 58B6 UBC10- 59NK- 60NK

Legend:
NEG (NTC) - Sample cancelled due to NTC Threshold.
NEG (R. Eff) - Sample cancelled as efficiency less than reaction efficiency threshold.

This report was generated by Rotor-Gene Q Series Software 2.3.1 (Build 49) Copyright 2013 QIAGEN GmbH. All Rights Reserved.