|  | www.qiagen.com |
| --- | --- |

# Quantitation Report

#### Experiment Information

| Run Name | Run 2021-12-15-sk4 |
| --- | --- |
| Run Start | 15.12.2021 13:53:27 |
| Run Finish | 15.12.2021 15:40:16 |
| Operator | MP |
| Notes |  |
| Run On Software Version | Rotor-Gene Q Software 2.0.2.4 |
| Run Signature | The Run Signature is valid. |
| Gain Green | 10, |
| Machine Serial No. | 0911113 |

#### Quantitation Information

| Threshold | 0,030 |
| --- | --- |
| Left Threshold | 5,000 |
| Standard Curve Imported | No |
| Standard Curve (1) | conc= 10^(-0,353*CT + 9,871) |
| Standard Curve (2) | CT = -2,835*log(conc) + 27,982 |
| Reaction efficiency (*) | (* = 10^(-1/m) - 1) 1,25309 |
| M | -2,83464 |
| B | 27,98183 |
| R Value | 0,98111 |
| R^2 Value | 0,96257 |
| Start normalising from cycle | 1 |
| Noise Slope Correction | Yes |
| No Template Control Threshold | % 0 |
| Reaction Efficiency Threshold | Disabled |
| Normalisation Method | Dynamic Tube Normalisation |
| Digital Filter | Light |
| Sample Page | Page 2 |
| Imported Analysis Settings |  |

#### Profile

| Cycle | Cycle Point |
| --- | --- |
| Hold 1 | Hold @ 95°C, 10min 0s |
| Cycling (40 repeats) | Step 1: Hold @ 95°C, 15s |
|  | Step 2: Hold @ 60°C, 30s, acquiring to Cycling A([Green][1][1]) |
|  | Step 3: Hold @ 72°C, 20s |
| Hold 2 | Hold @ 72°C, 1min 0s |
| Melt | Ramp from 55°C to 90°C |
|  | Hold for 90s on the 1st step |
|  | Hold for 5s on next steps,Melt A([Green][1][1]) |

#### Raw Data For Cycling A.Green

#### Quantitation data for Cycling A.Green

#### Standard Curve

| No. | Color | Name | Type | Ct | Ct Comment | Given Conc (IU/ml) | Calc Conc (IU/ml) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 21 |  | T6 1 1X | Standard | 17,79 |  | 10 000 | 3 926 |
| 22 |  | T6 1 1X | Standard | 17,50 |  | 10 000 | 4 966 |
| 23 |  | T6 1 10x | Standard | 18,72 |  | 1 000 | 1 852 |
| 24 |  | T6 1 10x | Standard | 18,90 |  | 1 000 | 1 601 |
| 25 |  | T6 1 100x | Standard | 21,33 |  | 100 | 221 |
| 26 |  | T6 1 100x | Standard | 21,41 |  | 100 | 208 |
| 27 |  | T6 1 1000x | Standard | 25,06 |  | 10 | 11 |
| 28 |  | T6 1 1000x | Standard | 24,96 |  | 10 | 12 |
| 29 |  | T6 1 10000x | Standard | 28,30 |  | 1 | 1 |
| 30 |  | T6 1 10000x | Standard | 29,14 |  | 1 |  |
| 31 |  | 1 NK | Unknown | 37,32 |  |  |  |
| 32 |  | 1 NK | Unknown | 36,51 |  |  |  |
| 33 |  | D6 1 ACO3 | Unknown | 19,99 |  |  | 658 |
| 34 |  | D6 1 ACO3 | Unknown | 16,83 |  |  | 8 608 |
| 35 |  | D6 1 ACO3 | Unknown | 15,93 |  |  | 17 872 |
| 36 |  | B6 1 ACO3 | Unknown | 16,78 |  |  | 8 935 |
| 37 |  | B6 1 ACO3 | Unknown | 17,34 |  |  | 5 676 |
| 38 |  | B6 1 ACO3 | Unknown | 17,60 |  |  | 4 615 |
| 39 |  | 1 NK | Unknown |  |  |  |  |
| 40 |  | 1 NK | Unknown |  |  |  |  |

-

Warning: The following samples were not analysed :
1T6 1 1X- 2T6 1 1X- 3T6 1 10x- 4T6 1 10x- 5T6 1 100x- 6T6 1 100x- 7T6 1 1000x- 8T6 1 1000x- 9T6 1 10000x- 10T6 1 10000x- 111 NK- 121 NK- 13D6 1 ACO1- 14D6 1 ACO1- 15D6 1 ACO1- 16B6 1 ACO1- 17B6 1 ACO1- 18B6 1 ACO1- 191 NK- 201 NK- 41T6 1 1X- 42T6 1 1X- 43T6 1 10x- 44T6 1 10x- 45T6 1 100x- 46T6 1 100x- 47T6 1 1000x- 48T6 1 1000x- 49T6 1 10000x- 50T6 1 10000x- 511 NK- 521 NK- 53D6 1 ACO3- 54D6 1 ACO3- 55D6 1 ACO3- 56B6 1 ACO3- 57B6 1 ACO3- 58B6 1 ACO3- 591 NK- 601 NK

Legend:
NEG (NTC) - Sample cancelled due to NTC Threshold.
NEG (R. Eff) - Sample cancelled as efficiency less than reaction efficiency threshold.

This report was generated by Rotor-Gene Q Series Software 2.3.1 (Build 49) Copyright 2013 QIAGEN GmbH. All Rights Reserved.