|  | www.qiagen.com |
| --- | --- |

# Quantitation Report

#### Experiment Information

| Run Name | Run 2021-12-15-sk6 |
| --- | --- |
| Run Start | 15.12.2021 13:53:27 |
| Run Finish | 15.12.2021 15:40:16 |
| Operator | MP |
| Notes |  |
| Run On Software Version | Rotor-Gene Q Software 2.0.2.4 |
| Run Signature | The Run Signature is valid. |
| Gain Green | 10, |
| Machine Serial No. | 0911113 |

#### Quantitation Information

| Threshold | 0,030 |
| --- | --- |
| Left Threshold | 5,000 |
| Standard Curve Imported | No |
| Standard Curve (1) | conc= 10^(-0,329*CT + 8,561) |
| Standard Curve (2) | CT = -3,037*log(conc) + 25,997 |
| Reaction efficiency (*) | (* = 10^(-1/m) - 1) 1,13449 |
| M | -3,0368 |
| B | 25,99718 |
| R Value | 0,97869 |
| R^2 Value | 0,95784 |
| Start normalising from cycle | 1 |
| Noise Slope Correction | Yes |
| No Template Control Threshold | % 0 |
| Reaction Efficiency Threshold | Disabled |
| Normalisation Method | Dynamic Tube Normalisation |
| Digital Filter | Light |
| Sample Page | Page 2 |
| Imported Analysis Settings |  |

#### Profile

| Cycle | Cycle Point |
| --- | --- |
| Hold 1 | Hold @ 95°C, 10min 0s |
| Cycling (40 repeats) | Step 1: Hold @ 95°C, 15s |
|  | Step 2: Hold @ 60°C, 30s, acquiring to Cycling A([Green][1][1]) |
|  | Step 3: Hold @ 72°C, 20s |
| Hold 2 | Hold @ 72°C, 1min 0s |
| Melt | Ramp from 55°C to 90°C |
|  | Hold for 90s on the 1st step |
|  | Hold for 5s on next steps,Melt A([Green][1][1]) |

#### Raw Data For Cycling A.Green

#### Quantitation data for Cycling A.Green

#### Standard Curve

| No. | Color | Name | Type | Ct | Ct Comment | Given Conc (IU/ml) | Calc Conc (IU/ml) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 41 |  | T6 1X | Standard | 14,70 |  | 10 000 | 5 237 |
| 42 |  | T6 1X | Standard | 14,68 |  | 10 000 | 5 348 |
| 43 |  | T6 10x | Standard | 15,28 |  | 1 000 | 3 394 |
| 44 |  | T6 10x | Standard | 17,16 |  | 1 000 | 811 |
| 45 |  | T6 100x | Standard | 20,08 |  | 100 | 89 |
| 46 |  | T6 100x | Standard | 19,34 |  | 100 | 156 |
| 47 |  | T6 1000x | Standard | 21,90 |  | 10 | 22 |
| 48 |  | T6 1000x | Standard | 22,18 |  | 10 | 18 |
| 49 |  | T6 10000x | Standard | 26,80 |  | 1 | 1 |
| 50 |  | T6 10000x | Standard | 27,13 |  | 1 |  |
| 51 |  | NK | Unknown |  |  |  |  |
| 52 |  | NK | Unknown |  |  |  |  |
| 53 |  | D6 UBC10 | Unknown | 11,94 |  |  | 42 535 |
| 54 |  | D6 UBC10 | Unknown | 12,01 |  |  | 40 427 |
| 55 |  | D6 UBC10 | Unknown | 11,85 |  |  | 45 548 |
| 56 |  | B6 UBC10 | Unknown | 12,70 |  |  | 23 920 |
| 57 |  | B6 UBC10 | Unknown | 12,65 |  |  | 24 923 |
| 58 |  | B6 UBC10 | Unknown | 12,60 |  |  | 25 838 |
| 59 |  | NK | Unknown | 16,16 |  |  | 1 736 |
| 60 |  | NK | Unknown | 21,80 |  |  | 24 |

-

Warning: The following samples were not analysed :
1T6 1X- 2T6 1X- 3T6 10x- 4T6 10x- 5T6 100x- 6T6 100x- 7T6 1000x- 8T6 1000x- 9T6 10000x- 10T6 10000x- 11NK- 12NK- 13D6 ACO1- 14D6 ACO1- 15D6 ACO1- 16B6 ACO1- 17B6 ACO1- 18B6 ACO1- 19NK- 20NK- 21T6 1X- 22T6 1X- 23T6 10x- 24T6 10x- 25T6 100x- 26T6 100x- 27T6 1000x- 28T6 1000x- 29T6 10000x- 30T6 10000x- 31NK- 32NK- 33D6 ACO5- 34D6 ACO5- 35D6 ACO5- 36B6 ACO5- 37B6 ACO5- 38B6 ACO5- 39NK- 40NK

Legend:
NEG (NTC) - Sample cancelled due to NTC Threshold.
NEG (R. Eff) - Sample cancelled as efficiency less than reaction efficiency threshold.

This report was generated by Rotor-Gene Q Series Software 2.3.1 (Build 49) Copyright 2013 QIAGEN GmbH. All Rights Reserved.