## Slide 1: Lasers - absorption methods
Vítězslav Otruba

## Slide 2: Lasers in AAS
Advantages against classical measuring radiation sources:
Narrow spectral line (linear calibration)
Continuous tuning of wavelength possible (scanning)
High intensity (low noise, fast processes - spark, furnace)
Low divergence (spatial profiles, mini atomizers)
Short pulse time (time resolution)
Any wavelength (AAS of excited and ionized atoms)

## Slide 3: AAS with low radiation intensity
Low intensity = not changed significantly occupation of energy levels
Atomic absorption profiles (Δλ≈ 0,1 – 0,001 pm)
Isotopic shifts (≈0,2 pm pro ΔM=1)
Splitting by nuclear spin
Doppler broadening (temperature)
Collision damping
Stark broadening (electron density)
Calibration graphs (linearity of up to 6 concentration orders)

## Slide 4: Absorption spectroscopy
Lambert's law:
Lambert-Beer law :
General equation for light absorption
Nn depends on the concentration of atoms (molecules) in a given state n.

## Slide 5: Absorption

## Slide 6: Linear absorption

## Slide 7: Fine structure Hg I 404,7 nm
Around 2/3 of the stable atomic nuclei has a resulting nuclear spin, causing a fine line structure.
199Hg and 201Hg have unpaired neutrons and show spin splitting. Resulting profile is then results of combination with the natural linewidth and Doppler broadening. Isotopes 198Hg; 200Hg; 202Hg and 204Hg are not splitted.

## Slide 8: AAS with graphite furnace
Example of AAS spectrometer with laser diode and electrothermal atomizer (graphite furnace)

## Slide 9: AAS in a laser spark
Spectral profile of the uranium absorption line. Distance to sample surface 5 mm, argon atmosphere 540 Pa.

## Slide 10: Molecular absorption spectroscopy
Identification of the composition of the sample is performed by monitoring the loss of light after passing through the sample
Laser excels extraordinary monochromaticity emitted radiation, wavelength selection is thus ensured directly by the radiation source
Laser radiation is extremely intense, so the signal to noise ratio is very high. Extremely low detection limits can be achieved when using a multireflective cuvette.
Measurement speed using a laser is higher than with conventional spectrometers. There is no parasitic infrared radiation from the heat source
The radiation source can be moved away from the measuring cuvette (using fiber optic)
Considerable attention is paid to laser spectroscopy in the IR region. Laser (diode) IR spectrometer (semiconductor tunable laser) allows very precise identification of spectral lines in IR (structural analysis)

## Slide 11: Absorption spectrometer with laser
A – classical absorption spectrometer; B – spectrometer with laser source

## Slide 12: Methods of low absorbance measurement
intense source of monochromatic light, absorption is measured differentially
multiple passage of absorbed light through the sample, increasing the absorption length
absorption of radiation and measurement of fluorescence from excited states
optoacoustic detection methods
absorption of radiation and measurement of Raman emission spectra

## Slide 13: Absorption spectroscopy in resonator cavity
Between the resonator mirror and the active laser medium is a space usable for spectroscopic purposes, in which a sample cuvette is inserted
This is a special detection method that excels in extraordinary sensitivity
Laser radiation passes through the cuvette located inside the resonator repeatedly
The laser output parameters are strongly influenced by internal absorption
Due to its sensitivity, the absorption method inside the resonator is best used to detect low concentrations of substances, especially gases
Spectroscopy inside the resonator is particularly suitable for qualitative analysis

## Slide 14: Absorption spectroscopy in resonator cavity
Active medium
Absorption cuvette
Pumping
Ф
100
10
1
10-9
10-7
10-5
g.cm-2 I
External cuvette
Internal cuvette
Absorption enhancement ≈ 5000x
Absorption environment: gas, flame,
plasma, solution…

## Slide 15: Differential measurement
Before the actual measurement it is necessary to set the gain of both photomultipliers so that the resulting voltage is zero. In the presence of plasma we then record the differential voltage in the order of magnitude V – mV.
monochromator
source
gray filter

## Slide 16: Multiple pass of light
basic disadvantages - we lose spatial resolution
- the light beam is diverging
source
detector

## Slide 17: Application of Fourier transform
The signal from the detector is modulated by the discharge. Fourier transform in a PC separates a signal with a different frequency than the modulation frequency; the result is noise reduction. It is also possible to realize direct modulation of the source (e.g. by lamp voltage). In absorption spectroscopy, it is suitable to use Fourier transform especially in the differential arrangement of detectors when the signal is small.
source
modulator
detector

## Slide 18: Cavity Ring Down Spectroscopy

## Slide 19: Cavity Ring Down Spectroscopy

## Slide 20: Cavity Ring Down Spectroscopy
Using memory oscilloscope and PC, a large number of pulses can be averaged and reducing noise.
The concentration of absorbing particles can be determined from the exponential decrease in light intensity.
The detection limit is of the order of A=10-6.
Parabolic mirrors allow spatial localization of the detected area, so concentration profiles can be measured.
Currently, this method is only used in the visible and near UV range.

## Slide 21: IR high resolution laser absorption spectrometry
FP = Fabry Perot standard for wave number calibration (Δν=c/(2dn))
L = semiconductor laser PbSnTe; PbCdS; Pb1-xSnxSe;
λ≈2 – 30 μm; R≈107;
UV – VIS: dye lasers, R≈5.107

## Slide 22: Double resonance methods
1
2
3
4
ω1,2
ω2,3
ω3,4
1
2 (virtual level)
3
ω1,2 = optical
ω2,3 = rf, mw
ω1,2 = constant frequency laser, strong population level 2
ω2,3 = tuned radio frequency, microwave or optical radiation
ω3,4 = fluorescence radiation indicating resonance

## Slide 23: Optical-microwave double resonance
detec-
tor
cuvette
laser active medium
mw
1
2
3
ω1,2
ω2,3
Indication: Absorption of auxiliary beam Pω1,2
At resonance 2 → 3, depopulation of level 2 occurs and its occupancy decreases, thereby increasing absorption 1 → 2
f≈ω2,3
f
A(Pω1,2)
Pω1,2
∼
detector
Pω1,2

## Slide 24: Tuning of absorption levels
Zeeman effect can be used for molecules with permanent magnetic dipole moment. The magnetic field causes splitting of degenerate levels:
ΔE = -μBgm, where g=Landé factor; μB=Bohr magneton; B= magnetic induction; m = magnetic quantum number.
Sensitivity is particularly high for radicals with unpaired electrons.
Stark effect causes the splitting of the molecules in a permanent electric dipole moment. Necessary intensities of homogeneous el. fields are of order 1000V/mm.
Especially in the IR region, lasers HF, DF, CO, CO2, N2O, H2O, D2O, HCN

## Slide 25: Tuning of absorption levels– Zeeman
detec-
tor
cuvette
lasers active medium
S
J
+3/2
+1/2
-1/2
-3/2
+1/2
-1/2
ΔM=-1
ΔM=0
ΔM=1
ΔM=1
B
→
E
B

## Slide 26: Nonlinear spectroscopic methods
Methods based on simultaneous absorption of multiple photons by the same sample particle
The absorption coefficient value changes when more photons are absorbed by the particles at the same time
Interaction of the sample with a large amount of photons increases the occupancy of the upper energy level and reduces the absorption of the sample due to saturation of the absorption transition
If non-linear effects occur, Lambert-Beer's law cannot be used for absorption

## Slide 27: Non-linear absorption

## Slide 28: Saturation spectrometry
Optically thick absorption - absorption coefficient regulation
Spectroscopy without Doppler Broadening e.g.:
  Lamb-dip spectrometry
  Two-photon Doppler-free spectrometry
Laser frequency stabilization
Multiphoton methods

## Slide 29: Saturation Doppler-free Spectroscopy (Lamb-Dip Spectroscopy)
The saturation sub Doppler spectroscopy method finds its main application in determining the exact values of absorption lines and in stabilizing of lasers
The principle of saturation sub Doppler spectroscopy is based on the Doppler effect
Saturation sub Doppler spectroscopy is a method used for the study of gaseous substances
Gas particles that move chaotically appear in radiation interaction with frequency shift according to the speed of movement relative to the viewing direction
The particles interact with the frequency shift radiation given by the instantaneous velocity component

## Slide 30: Natural linewidth

## Slide 31: Heisenberg uncertainty principle

## Slide 32: Doppler broadening

## Slide 33: Homogeneous and inhomogeneous broadening

## Slide 34: Movement of particles in gas
Gas particles that move chaotically appear in interaction with radiation by frequency shift according to the speed of movement relative to the viewing direction
hν
v0-vD v0 v0+vD

## Slide 35: Absorption saturation on inhomogeneously broadened line

## Slide 36: Both the saturation and probe beams are parallel
λ0
λs=λm=λ0+λD
λs
λm
D
D
Doppler
Lorentz
λ0; λ0 ± λD

## Slide 37: Bennet (Lamb) hole

## Slide 38: Saturation and probing beam in opposite
λ0
λs
Lorentz
λm ≈ λ0-λD
λs
λm
D
D
Doppler
λ0; λ0 ± λD

## Slide 39: Final solution
λs
D
D
λ0; λ0 ± λD
Doppler
λs=λm=λ0
Lorentz

## Slide 40: Saturated absorption spectroscopy

## Slide 41: Doppler-free spectroscopy

## Slide 42: Experimental arrangement
Experiment results:
  Na(g),p=40μPa,
  t=110°C, λ=589 nm
  ΔλDOP ≅1,7 pm
  (1500 MHz)
  ΔλLAS ≅0,008 pm
  (7 MHz)
  ΔλEXP≅0,068 pm
  (40 MHz)
  ΔλNAT ≅0,01 pm (≈τ=16 ns)

## Slide 43: Apparatus for Doppler-free saturated absorption spectroscopy of I2 and Na

## Slide 44: Doppler-Free Saturated Absorption Spectroscopy of Iodine and Sodium Using a Tunable Ring Dye Laser

## Slide 45: Absorption spectrum of iodine
Ordinary, Doppler broadened, (dashed line) and (b) Doppler-free (solid line)
absorption spectra of the 5682 Å, P(117), 21-1, X --> B transition of 127I2.
(c) 300 MHz interferometer transmission peaks for frequency calibration of laser scan.

## Slide 46: Laser stabilization
L – laser active medium, Z – mirrors, K = cuvette,
E – etalon, PP = piezoelectric interface, LO – tuning circuit,
D – detector, R – feedback control.
Example: He-Ne laser 3390 nm, methane cuvette, stabilization ±0,5 Hz.

## Slide 47: Doppler-free spectroscopy - summary
If the frequency of the intense laser beam is tuned off the centre of the absorption line profile, a dip is formed in the absorption line indicating a decrease in the absorption coefficient (non-linear effect)
In the experiment the laser beam is divided into two beams, intensive - saturation and weaker - test beams
Both beams pass through the cuvette against each other - the effects of the two beams, when aligned from the center of the absorption line, are positioned symmetrically with respect to the center of the line
When tuning the laser to the center of the line, the test beam is less absorbed due to the decrease in absorption in the dip and the laser beam intensity is significantly increased
A narrow resonance peak is obtained which represents the absorption line of the sample without Doppler broadening
Doppler-free spectroscopy allows experimentally removing the line broadening by selecting particles of zero velocity from the sample

## Slide 48: Experimental results

## Slide 49: Two-photon Doppler-free spectroscopy
Δω=ωL(v/c)
Δω=-ωL(v/c)
ωL
ωL
0
V
1
Δω=0
Δω = Doppler shift
τv ≈10-12 s
ωL =½ E01
VIS: Δλ≈1 GHz
ФL<dyn. Stark efect
ΔλL ≈ 1 MHz ⇒
necessary stabilization by combination of Zeeman and Stark tuning

## Slide 50: Experimental arrangement
Laser
Detector
absorption
Cuvette
Detector
fluorescence
A
FL
Detection of fluorescence in UV-VIS and near IR - transitions from excited levels
Detection of absorption of one of the exciting beams mainly in the IR region, e.g. UV-VIS

## Slide 51: Two-photon spectrometry

## Slide 52: Application