## Slide 1
How the NMR Spectrometer Works
Magnet

## Slide 2
How the NMR Spectrometer Works
Magnet – homogeneity adjustments (shims)

## Slide 3
How the NMR Spectrometer Works
Probe
18.10.2011

## Slide 4
How the NMR Spectrometer Works
!!!!!!!!B1 ~ I ~ P1/2 (P=RI2)!!!!!!!!
Lp =
[dB]
Transmitter – power vs. rf field induction
Relative power ratio expressed in decibel
1 dB: P1/P2 = 1.2589254
Relative rf field ratio ex-pressed in decibel
1 dB: w1/w2 = 1.120185

## Slide 5
How the NMR Spectrometer Works
Transmitter – phase shifted pulses

## Slide 6
How the NMR Spectrometer Works
A/D converter
n=6 bits
64 levels
n=8 bits
256 levels
3bit A/D converter (8 levels)
2n = 8
A/D converter typically 16 bits
i.e. 65 536 levels
32 bits
i.e. 4 294 967 296 levels

## Slide 7
How the NMR Spectrometer Works
A/D converter – sampling frequency
Nyquist frequency
fmax = 1 kHz => D = 500 ms

## Slide 8
How the NMR Spectrometer Works
A/D-converter – signal folding

## Slide 9
How the NMR Spectrometer Works
Receiver
NMR signal
Local oscilator
Intermediate frequency
A/D converters

## Slide 10
How the NMR Spectrometer Works
Quadrature detection
mixer
cosA.cosB=1/2(cos(A+B) + cos(A-B))
cosA.sinB=1/2(sin(A+B) - sin(A-B))
(
)

## Slide 11
How the NMR Spectrometer Works
Quadrature detection
Intermediate frequence
Frequency wo+wrx removed by filter

## Slide 12
How the NMR Spectrometer Works
Quadrature detection – time vs. frequency
Spectral width fsw
Acqusition time tacq
N – number of acquisition points
D – sampling interval

## Slide 13
How the NMR Spectrometer Works
Pulse programmer
INEPT with refocusing

## Slide 14
How the NMR Spectrometer Works
Pulse programmer
;ineptrd
;avance-version (02/05/31)
;INEPT for non-selective polarization transfer
;with decoupling during acquisition
#include <Avance.incl>
"p2=p1*2"
"p4=p3*2"
"d3=1s/(cnst2*cnst11)"
"d4=1s/(cnst2*4)"
"d12=20u"
1 ze
2 30m do:f2
d1
d12 pl2:f2
(p3 ph1):f2
d4
(center (p4 ph2):f2 (p2 ph4) )
d4
(p3 ph3):f2 (p1 ph5)
d3
(center (p4 ph2):f2 (p2 ph6) )
d3 pl12:f2
go=2 ph31 cpd2:f2
30m do:f2 mc #0 to 2 F0(zd)
exit
ph1=0 0 0 0 0 0 0 0 2 2 2 2 2 2 2 2
ph2=0 2
ph3=1 1 3 3
ph4=0 2
ph5=0 0 0 0 1 1 1 1 2 2 2 2 3 3 3 3
ph6=0 2 0 2 1 3 1 3
ph31=0 0 2 2 1 1 3 3
;pl1 : f1 channel - power level for pulse (default)
;pl2 : f2 channel - power level for pulse (default)
;pl12: f2 channel - power level for CPD/BB decoupling
;p1 : f1 channel - 90 degree high power pulse
;p2 : f1 channel - 180 degree high power pulse
;p3 : f2 channel - 90 degree high power pulse
;p4 : f2 channel - 180 degree high power pulse
;d1 : relaxation delay; 1-5 * T1
;d3 : 1/(6J(XH)) XH, XH2, XH3 positive
; 1/(4J(XH)) XH only
; 1/(3J(XH)) XH, XH3 positive, XH2 negative
;d4 : 1/(4J(XH))
;d12: delay for power switching [20 usec]
;cnst2: = J(XH)
;cnst11: 6 XH, XH2, XH3 positive
; 4 XH only
; 3 XH, XH3 positive, XH2 negative
;NS: 4 * n, total number of scans: NS * TD0
;DS: 16
;cpd2: decoupling according to sequence defined by cpdprg2
;pcpd2: f2 channel - 90 degree pulse for decoupling sequence
;$Id: ineptrd,v 1.8 2002/06/12 09:05:00 ber Exp $
23.10.2012