## Slide 1: NMR Sample

## Slide 2: Sample quantity
NMR has intrinsic­ally poor sensitivity
Concentration
  Simple spectra (proton 1D) ~ 50 μM
  Multidimensional spectra ~0,5 mM
Sample tubes
  5 mm, 550 μl
  3 mm, 150 μl
  Shigemi 5 mm, 250 μl

## Slide 3: Solubility and stability
Aggregation, precipitation
NA better soluble than proteins
NMR measurements long
  Run a control spectrum before and after a long experiment (1D, 1H-15N HSQC for labeled proteins)

## Slide 4: Molecular weight
Increasing molecular weight
  Complexity of the spectrum – number of lines, overlap
  Increasing linewidths
Remedies
  Isotope labeling
  Multidimensional spectra
  TROSY, MQ

## Slide 5: NMR tube
Choice
  Volume
  Quality (geometry)
Cleaning
  High quality tubes are expensive, not for single use!
Drying
  No high temperatures above 50⁰C

## Slide 6: Concentration, solvent exchange
Lyophylization
  May harm samples (esp. proteins)
Ultrafiltration
Gel filtration
Ultracentrifugation

## Slide 7: Isotope effect, H2O vs D2O
No D2O is 100%
Viscosity
Melting/freezing point of D2O is 3.8⁰C
pH
  Glass electrode shows 0.4 units difference (7.0 in H2O, 6.6 in D2O), pH* value

## Slide 8: Sample parameters
pH
Ionic strength
  NMR probes (esp. cryoprobes) do not like high salt concentrations (above 100 mM)
Buffer type
  Preferably buffer with no NMR signals
Concentration
  Oligomerization, precipitation, high viscosity
Temperature
  Sharper lines at higher temperatures, but not for exchangeable protons

## Slide 9: Chemical shift reference
Internal
  Add reference compound to the sample
External
  Capillary within the sample tube
Direct
  Reference compound measured together with the sample
Indirect
  Reference compound measured separately or the chemical shift is calculated

## Slide 10: Common reference compounds Wishart et al. (1995) , J. Biomol. NMR 6, 135-140.
1H
  TMS 0.00 ppm, not soluble in water
  TSP around 0, varies slightly with pH
  DSS 0.00 ppm
13C
  TMS 1.7 ppm, not soluble in water
  TSP around 0, varies slightly with pH
  DSS 0.00
15N
  NH3 0.00 ppm, bio
  CH3NO2 (neat) 381.7, chemistry
31P
  H3PO4
TMS
tetramethylsilane
DSS
Sodium trimethylsilylpropanesulfonate
(dimethyl silapentane sodium sulfonate)
TSP
Trimethylsilylpropanoic acid

## Slide 11: Contamination
Paramagnetic contamination
  Broadens the lines, destroys resolution
  Cu ,Mn, Cr, Fe, Co
Microbial contamination
  Bacteria can eat-up your sample
  Work clean, add sodium azide (toxic!), EDTA
  Bacteria do not strive in 100% D2O