## Slide 1: 1 Introduction
Avance 1D/2D
BRUKER
9
This manual gives an introduction into basic one- and two-dimensional nuclear magnetic resonance (NMR) spectroscopy. After a short introduction the acquisition of basic 1D 1H and 13C NMR spectra is described in the Chapters 3 to 7. Homonuclear 2D [1H,1H] correlation spectra are described in Chapter 8 (COSY), 9 (TOCSY), 10 (ROESY) and 11 (NOESY).
Heteronuclear 2D [13C,1H] correlation experiments are described in Chapter 12 (XHCORR), 13 (COLOC), 14 (HMQC) and 15 (HMBC). The Chapter 16
contains the description of inverse 2D [13C,1H] correlation experiments using pulsed field gradients, and some special NMR experiments are described in
chapters 17 to 20. A brief introduction to NMR automation with the IconNMR program is given in chapter 21.
1.1 An Important Note on Power Levels
Several times throughout this manual, the user is asked to set the power levels pl1, pl3, etc. to the “high power” level for the corresponding channel (f1 or f2). In order to avoid damaging the probehead or other hardware components, the user is advised to use only the power levels indicated in Table 1 below, if no other information (e.g. final acceptance tests) is available.
Note that these “power levels” are really attenuation levels, and so a higher value corresponds to a lower power. Also note that these power levels pertain only to the specific spectrometers and amplifiers listed below, which correspond to the AVANCE instruments as of July 2000. It is assumed that no correction tables (CORTAB) are existing.
Table 1: Suggested “Proton and Carbon High Power” Levels for Avance Instruments
Nucleus | Spectrometer | Amplifier | Power Level
1H | Avance | BLA2BB |  + 3dB
 |  | BLARH100 |  + 3dB
 |  | BLAXH300/50 |  0dB
 | Avance DPX | BLAXH20 | = - 6dB
 |  | BLAXH40 | = - 3dB
 |  | BLAXH100/50 |  0dB
 | Avance DRX | BLAXH150/50 |  0dB
 |  | BLAXH300/50 |  0dB
 | Avance DMX | BLARH100 |  + 3dB

## Slide 2
10
BRUKER
Avance 1D/2D
Nucleus | Spectrometer | Amplifier | Power Level
 | Avance | BLA2BB |  + 6dB
 |  | BLAX300/50 |  + 6dB
 |  | BLAX300 |  + 6dB
 |  | BLAX500 |  + 9dB
 | Avance DPX | BLAXH20 | = - 6dB
13C |  | BLAXH40 | = - 6dB
 |  | BLAXH100/50 |  - 3dB
 | Avance DRX | BLAXH40 |  - 3dB
 |  | BLAXH150/50 |  0dB
 |  | BLAXH300/50 |  6dB
 | Avance DMX | BLAX300 |  + 6dB
 |  | BLAX500 |  + 9dB
1.2 NMR Spectrometer
The NMR spectrometer consists of three major components: (1) The superconducting magnet with the probe, which contains the sample to be measured; (2) The console, which contains all the electronics used for transmission and reception of radio frequency (rf) pulses through the pre- amplifier to the probe; (3) The computer, from where the operator runs the experiments and processes the acquired NMR data.
1.3 Classical Description of NMR
A more complete theoretical description of NMR is given in chapter 22.
Among the various atomic nuclei, about a hundred isotopes possess an intrinsic angular momentum, called spin and written ħI . They also possess a magnetic moment  which is proportional to their angular momentum:
  ħI
where  is the gyromagnetic ratio.
The Larmor theorem states that the motion of a magnetic moment in a magnetic field B0 is a precession around that field, where the precession frequency is given by:
  B Larmor frequency
0 0

## Slide 3
Avance 1D/2D
BRUKER
11
By convention, the external static field (B0) is assumed to be along the z-axis and the transmitter/receiver coil along either the x- or y-axis. After the sample has been inserted into the magnetic field it shows a magnetization vector M along the z-axis. In this state, no NMR signal is observed, as we have no tranverse rotating magnetization.
By application of an additional rotating magnetic field B1 in the x-y-plane, the orientation of M can be tilted into the x-y plane where it precesses around the total magnetic field, e.g. the vector sum of B0 and B1. Such a rotating magnetic field is obtained by applying rf-pulses, and the components of M are described by the Bloch equations:
dt
d M
dt
z y 1
y 1 z
 M rB
d M r  B M
dt
x
d M r  0
Assuming the magnetization at time 0 to be along the z-axis with amplitude
M0, we find the following solution to the above equation:
M z (t)  M 0 cos(B1t)
The magnetization vector is precessing around the B1 axis which is aligned with the x-axis of the reference system. If we choose the time t of suitable duration, we obtain:
r
y 0 1
M (t)  M sin(B t)
2
1
  B t  
which is defined as the 90 degree pulse creating maximum y-magnetization, which in turn yields maximal signal intensity.
1.4 Spin Operators of a One-Spin System
All NMR experiments start from the thermal equilibrium. In thermal equilibrium, the classical description gives a magnetic moment parallel to the static field, Mz. In the Spin Operator formalism, this is described by:
  I
eq z
where eq is the equilibrium density matrix representing the state of the spin system under investigation.
Now there are only two basic types of evolutions: (1) An external perturbation, e.g. a rf-pulse, or (2) an unperturbed evolution which will eventually bring the system back to the thermal equilibrium.

## Slide 4
12
BRUKER
Avance 1D/2D
1.4.1 Effect of rf-Pulses
The effect of an rf-pulse is that of a rotation along the pulse axes according to the following calculus rules:
x
Iz  Iz cos   I y sin 
 y
Iz  I z cos   Ix sin 
 x
Ix  I x
 y
I y  I y
 y
Ix  Ix cos   I z sin 
 x
I y  I y cos   Iz sin 
If the flip angle  = 90° then:
Iz  Ix, y
90 y ,x
90 y,x
Ix, y  Iz
We find the expected result, that a 90° pulse will generate transverse magnetization. The rest of this chapter will be concerned with finding out about the fate of this transverse magnetization in time.
We introduced tacitly the arrow notation, where we find on the left side the system before and on the right side after the specific evolution under the operator noted above the arrow. This notation is simple, very convenient and not only limited to the description of rf-pulses. We will discuss this notation in more details in the next section.
1.4.2 Effect of Chemical Shift Evolution
The so-called chemical shift Hamiltonian is given by:
H    I z
where  is the chemical shift of the corresponding nucleus in the NMR spectrum ( 0  where 0 is the Larmor frequency of the spin and  the carrier frequency of the interaction frame).
The calculus rules for the chemical shift evolution are the following:
I Iz t I
z z
I Iz  t I cos( t)  I sin( t)
x x y
I Iz t I cos( t )  I sin( t)
y y x
The time t is the period, during which the Hamiltonian is valid. The Hamiltonian of a spin system can change with time, for example if the experimental setup prescribes first a rf-pulse and then a period of unperturbed evolution. For the calculus rules it is mandatory, that each Hamiltonian is time independent during the time t.
What’s the general idea? The whole NMR experiment is divided into time intervals, during which the Hamiltonian can be made time independent by

## Slide 5
Avance 1D/2D
BRUKER
13
choice of a suitable interaction frame. Typical experiments are divided in pulse intervals and free evolution times.
During the pulses, the chemical shift and scalar coupling interaction is ignored. Only the applied B1 field is considered. This approach is justified for pulses with tpulse«T1,T2.
1.4.3 Effect of Scalar Coupling
Apart from the chemical shift, there is a second very import interaction between spins, the scalar coupling. The scalar depends on the mediation of electrons, which are confined in orbitals around both nuclei. The scalar coupling is expressed in Hz and noted as J. The operator expression for the scalar coupling is:
2 J12 I1zI2z
The above Hamiltonian expresses the scalar coupling between spin 1 and spin 2 with a coupling constant J12. The evolution Hamiltonian for this spin system is then:
H  1 I1z  2 I2 z  2 J12 I1z I2z
To calculate the effect of this Hamiltonian, it is divided into 3 parts:
1 I1z
2 I2z
2 J12 I1zI2z
which are applied in sequence, where this sequence is arbitrary. After a 90° pulse has been applied to the two spins, we first calculate the two chemical shift terms:
2 I2z t  I cos( t)  I sin( t)
1x 1 1y 1
 I2x cos(2 t)  I2 y sin(2 t)  1
The next step will be to compute the evolution under the scalar coupling. The scalar coupling term can be evaluated with a simple set of rules:
I 2 J12 I1zI2z t I
1z 1z
I 2 J12 I1zI2z t I cos(J t)  2 I I sin(J t)
1x 1x 12 1y 2z 12
I 2 J12 I1zI2z t I cos(J t)  2 I I sin(J t)
1y 1y 12 1x 2 z 12
2  I I 2 J12 I1zI2z t 2 I I cos(J t)  I sin(J t)
1x 2z 1x 2 z 12 1y 12
2  I I 2 J12 I1zI2z t 2 I I cos(J t)  I sin(J t)
1y 2z 1y 2z 12 1x 12
2  I I 2 J12 I1zI2z t 2 I I
1x 2y 1x 2 y
which can then be applied to the various terms of 1 above:
2 z
1x
eq 1z 2 z
cos(1 t)  I1y sin(1 t)  I
  I  I
1 I1z t  I

## Slide 6
14
BRUKER
Avance 1D/2D
 2 J12 I1zI2z t {I cos(J t)  2 I I sin(J t)} cos( t)
1 1x 12 1y 2z 12 1
{I cos(J t)  2 I I sin(J t)} sin( t)
1y 12 1x 2z 12 1
{I cos(J t)  2 I I sin(J t)} cos( t)
2 x 12 1z 2 y 12 2
{I cos(J t)  2 I I sin(J t)} sin( t)
2 y 12 1z 2 x 12 2
 
2
References: O. W. Sørensen, G.W. Eich, M. H. Levitt, G. Bodenhausen, R.
R. Ernst, Progres in NMR Spectroscopy, 16, 163 (1983).
1.5 Sensitivity of NMR Experiments
The sensitivity of NMR experiments is given by the signal to noise ratio:
ns
B 3 / 2
N T 
S / N  exc 2 det 0
T
S/N | = | signal to noise ratio
N | = | number of spins in the system (sample concentration)
exc | = | gyromagnetic ratio of the excited nucleus
det | = | gyromagnetic ratio of the detected nucleus
ns | = | number of scans
B0 | = | external magnetic field
T2 | = | transverse relaxation time (determines the line width)
T | = | sample temperature
(Comment: here we can already see that it might be useful for a better signal to noise ratio to excite one kind of nuclei and detect another kind with a better gyromagnetic ratio in the same experiment. This is done in inverse experiments which are described in sections 14 to 16).
  Useful Coupling Constants
Many NMR constants such as chemical shift ranges, sensitivities, common NMR solvent properties etc. can be found in the Bruker Almanac. Here we added the values of some common coupling constants that are used more often as parameters (cnst1 – cnst5) in some pulse programs.
    Coupling Constants: nJCH
As a rule of thumb it is possible to estimate the 1JCH coupling constant from the following equation: 1JCH ~ 500*(fractional CH s character). That is: 125Hz
< 1JCH < 250Hz, so that 1JCH = 145Hz is a good approximation in most cases.
The values of 2JCH coupling constants increase with increasing HCC angles and with the electronegativity of the C substituent They vary between –5 and 50Hz.

## Slide 7
Avance 1D/2D
BRUKER
15
The 3JCH coupling constants are mostly positive and are maximal at CCCH angles of 0° and 180°. The values for trans couplings are larger as for cis couplings (Karplus relation).
Table 2: Useful CH Coupling Constants
Compound | 1JCH in Hz
Ethane | 124.9
Acetonitrile | 136.0
Ethene | 156
Benzene | 159
Dichloromethane | 178.0
Chloroform | 209.0
Formaldehyde | 222.0
System | 2JCH in Hz
C(sp3)C(sp3)H | -10 to +6
C(sp3)C(sp2)H | 0 to +30
C(sp2)C(sp3)H | -7 to -4
C(sp2)C(sp2)H | -4 to +14
System | 3JCH in Hz
C(sp3)C(sp3)C(sp3)H | 0 to 8
C(sp3)C(sp2) C(sp2)H | 0 to 20
C(sp2)C(sp2) C(sp3)H | 0 to 20
 | Jtrans > Jcis
References: H.-O. Kalinowski, S. Berger, S. Braun; 13C-NMR-Spektroskopie;
Georg Thieme Verlag; Stuttgart, New York.
1.6.2 Coupling Constants of Hydrocarbons: nJHH
Usually 2JHH coupling constants are negative and vary in a range between
3JHH
-0.5Hz and -15Hz in hydrocarbons. coupling constants are mostly
positive and usually range from 2 up to 18Hz. The n>3JHH coupling is positive or negative with smaller absolute values, that range from 0 to 3Hz. The Karplus relation is also valid: Jtrans > Jcis.
Table 3: Useful HH Coupling Constants
System | 2JHH in Hz
HC(sp3)H | -12 to -15
HC(sp2)H | -0.5 to -3
System | 3JHH in Hz
HC(sp3)C(sp3)H | 2 to 9
HC(sp3)C(sp2)H | 4 to 10
HC(sp2)C(sp2)H | 6 to 18
HC(sp3)CHO | 1 to 3
HC(sp2)CHO | 2 to 4
System | 4JHH (abs. value) in Hz
HC(sp3)C(sp3)C(sp3)H | 0
HC(sp3)C(sp2)C(sp2)H | 0 to 3
HC(sp)C(sp)C(sp3)H | 2 to 3

## Slide 8
16
BRUKER
Avance 1D/2D
Heteroatoms with considerable I or M effect can shift the J values dramatically.

## Slide 9: 2 Preparing for Acquisition
Avance 1D/2D
BRUKER
17
  Sample Preparation
The sample quality can have a significant impact on the quality of the NMR spectrum. The following is a brief list of suggestions to ensure high sample quality:
    Always use clean and dry sample tubes to avoid contamination of the sample.
    Always use high quality sample tubes to avoid difficulties with shimming.
    Filter the sample solution.
    Always use the same sample volume or solution height (recommended values: 0.6 ml or 4 cm of solution for 5 mm sample tubes, 4.0 ml or 4 cm of solution for 10 mm sample tubes). This minimizes the shimming that needs to be done between sample changes.
    Use the depth gauge to position the sample tube in the spinner. This is discussed further in Chapter 5 ‘Sample Positioning' of the BSMS User's Manual.
    Check that the sample tube is held tightly in the spinner so that it does not slip during an experiment.
    Wipe the sample tube clean before inserting it into the magnet.
    For experiments using sample spinning, be sure that the spinner, especially the reflectors, are clean. This is important for maintaining the correct spinning rate.
  Bruker NMR software
There are three major tasks that are controlled by the NMR software: acquisition, processing and plotting. The XWinNMR program is the user interface for all of these tasks. The commands can either be called up by selecting the menu items or by typing the appropriate command in the command line followed by RETURN. There are many parameters that are important for each job and they can be accessed and edited by the user. These parameters and the measured data as well as the processed spectra are stored in datasets which are specified by names, experiment numbers (expno) and processing numbers (procno).

## Slide 10
18
BRUKER
Avance 1D/2D
Each parameter can be accessed directly by entering it’s name in the command line followed by RETURN or in the eda, edp or edg window for acquisition-, processing- or plotting parameters respectively. Since these panels contain all possible parameters and are rather large, it is often more convenient to use somewhat more reduced parameter editor interfaces. The
ased command opens the panel for the acquisition parameters that are of
importance only for the selected pulse program. Here the parameters are also commented on.
2.2.1 Predefined Parameter Sets
The XWinNMR philosophy is to work with predefined parameter sets that are suitable for most of the NMR tasks and experiments you are facing. These parameter sets include the pulse program, acquisition and processing AU programs as well as all other necessary parameters except spectrometer specific values for pulse lengths and power levels. These standard parameter sets usually have the same base name as the corresponding pulse program. Each parameter set can be called up into a dataset of your choice by the command rpar. You can modify the parameters and save the new parameter set by the command wpar. Bruker predefined parameter sets are written in capital letters, and we recommend that you do not change them but rather create new ones that you can use just as well.
Therefore the most simple way to run a certain experiment is to create a new dataset with a specific name, using the command edc. Then you would read the corresponding parameter set by rpar (i.e. rpar PROTON all), set the pulse lengths and power levels by getprosol and type xaua to start the acquisition. (It is assumed that the sample is shimmed and the probe is matched and tuned for the specific nuclei). If you are using the Bruker predefined parameter sets, you can always process the data by typing xaup.
The following list is a short summary of the most commonly used experiments and the corresponding parameter sets. The emphasis is on the spectroscopic information that you will get from the experiments rather than on the type of experiment. (For the experiments in this table, it is always recommended to use the gradient version of the experiment if you have the required z-gradient hardware. These experiments usually require less time than the ones without gradients).
Table 4: Short List of Typical Experiments, Parameter Sets and What They Do

## Slide 11
Avance 1D/2D
BRUKER
19
Atom / Group | Information (1D Experiments) | a.k.a. | Parameter Set
H | 1H chemical shift and coupling | 1D1H | PROTON
C | 13C chemical shift, 1H decoupled (signal enhancement, integration not possible) | 1D13C | C13CPD
C | 13C chemical shift, 1H coupled (signal enhancement, integration not possible) | 1D13C | C13GD
CH, CH2, CH3 | 13C chemical shift, select CH, CH2 and CH3 signals only (same phase) | DEPT45 | C13DEPT45
CH | 13C chemical shift, select CH signals only | DEPT90 | C13DEPT90
CH, CH3 | 13C chemical shift, select CH and CH3 signals only (opposite phase) | DEPT135 | C13DEPT135
Correlation | Information (2D Experiments) | a.k.a. | Parameter Set
H–H | 1H/1H nearest neighbor, through bond chemical shift correlation | COSY | COSYGPSW1 COSY45SW
H–H | 1H/1H nearest neighbor, through bond chemical shift correlation plus coupling constants | DQF- COSY | COSYGPDFPHSW1 COSYDQFPHSW
H–(–)nH | 1H/1H total spin system through bond chemical shift correlation | TOCSY | MLEVPHSW
C–H | Sensitive 1H/13C directly bound chemical shift correlation (one bond), lower resolution in 13C dimension | HSQC HMQC | HSQCGPPH1 HMQC
C–H | Sensitive 1H/13C directly bound chemical shift correlation (one bond), lower resolution in 13C dimension (small molecules, solemnly select 13C/1H not 12C/1H) | BIRD- HMQC | HMQCBI
C–H | Insensitive 1H/13C directly bound chemical shift correlation (one bond), high resolution in 13C dimension | HETCOR | HCCOSW
C–(–)nH | Sensitive 1H/13C long range chemical shift correlation (more than one bond), lower resolution in 13C dimension | HMBC | HMBCGPLPND1 HMBCLPND
C–(–)nH | Insensitive 1H/13C long range chemical shift correlation (one and more bonds), high resolution in 13C dimension | COLOC | HCCOLOCSW
H…H | 1H/1H non bound structural neighbor, through space chemical shift correlation (small molecules, low fields) | ROESY | ROESYPHSW
H…H | 1H/1H non bound structural neighbor, through space chemical shift correlation (large molecules, proteins) | NOESY | NOESYPHSW
In most of the 2D parameter sets there is a spectral width optimization implemented (PULSEPROGRAMSW). So if you acquire the corresponding 1D experiments in the previous experiment number the spectral width for the 2D will be optimized according to the 1D information.
A complete list of parameter sets can be called up by typing rpar without a following name. The nomenclature of the parameter sets follows the rules for
1 z-gradient hardware required

## Slide 12
20
BRUKER
Avance 1D/2D
the nomenclature of the pulse programs. They can be found in the file:
$XWinNMRHome/exp/stan/nmr/lists/pp/Pulprog.info
However in this manual, we focus on the manual setup of the experiments from scratch and the optimization of the corresponding parameters. therefore the rpar command will not be used throughout this text.
2.2.2 XWinNMR parameters and commands
A list of commonly used acquisition and processing commands and parameter names as well as a description of the corresponding command or parameter is given in short in the tables below.
Table 5: General Commands and AU Programs
setres | customize the XwinNMR interface
edmac | edit or create an XWinNMR macro
edau | edit or create an XWinNMR AU program
edpul | edit or create an XWinNMR pulse program
xau listall_au | create a file called “listall” in your home directory with a list of all available AU programs including short descriptions
edcpul | edit the current pulse program
Table 6: Data Set Related Commands
edc, new | create a new data set, experiment number or processing number
xau iexpno | copy the current experiment number including all parameters to the consecutive experiment number
wrpa | copy of the current data set including the spectra
re | move to a specific experiment number within the data set
rep | move to a processing number within the experiment number
browse | browse the data set directories
search | find a specific data set
wpar | save the current parameters
rpar | select and read a predefined parameter set
Table 7: Acquisition Parameters
ns | number of scans
ds | number of dummy scans
td | Time domain, number of acquired data points
sw | sweep width in ppm
aq | acquisition time
o1p | transmitter frequency of f1 channel in ppm
o2p | transmitter frequency of f2 channel in ppm
rg | receiver gain
pulprog | definition of the pulse program
aunmp | definition of the acquisition AU program
Table 8: Acquisition and Pre-acquisition Commands
edhead | define the current probehead
edprosol | define probehead specific pulse lengths and power levels
getprosol | use probehead specific pulse lengths and power levels in the

## Slide 13
Avance 1D/2D
BRUKER
21
 | current pulse program
xau pulse | calculate the power level from pulse lengths and vice versa
edsp, edasp | configure the routing of the spectrometer
edcpul | open the current pulse program in a text editor window
eda | edit all acquisition parameters
ased, as | edit the acquisition parameters that are relevant for the current pulse program
ppg | graphical display of the current pulse program
spdisp | open the graphical pulse program editor
dpa | display all status parameters for the acquisition
wbchan | select the wobbling channel for tuning and matching
wobb | tuning and matching the probe
atma | automatic tune and match the ATM probe
atmm | manually tune and match the ATM probe
edsolv | define solvent parameters
edlock | define lock parameters for probhead and solvent
lock | Automatically lock on solvent (parameters defined in edlock)
lockdisp | open the lock display window
rsh | select and read shim values
gradshim | start the gradient shimming subprogram
wsh | save the current shim values
edte | open the temperature control window
edau | select or edit AU programs
stdisp | open the shape tool
expt | estimate the experiment time
rga | Automatically adjust the receiver gain
zg | start acquisition
xaua | start the acquisition AU program (this also starts the acquisition)
gs | Interactive adjustment of acquisition parameters
tr | data transfer during acquisition
halt, stop | stop the acquisition
kill | kill a specific process
Table 9: Processing Parameters
si | size of the real spectrum
phc0, phc1 | Parameters for zero order and first order phase corrections
lb | line broadening factor for em
aunmp | definition of the processing AU program
Table 10: Processing Commands
edp | edit all processing parameters
dpp | display all status parameters for processing
ft | Fourier transform the current data
em | apply exponential window function
ef | combined command of ft and em
phase | set the phase correction defined by phc0 and phc1
apk | Automatically phase correct the spectrum
abs | Automatically baseline correct and integrate the spectrum
efp | combined command of ft, em and phase

## Slide 14
22
BRUKER
Avance 1D/2D
sr | spectral referencing
sref | Automatically calibrate the spectrum
edc2 | select a second and a third data processing number
dual | invoke the dual display
edo | select an output device
edg | edit all graphics and plotting parameters
view | plot preview
xwinplot | start the plot program
Table 11: Pulse Program Specific Parameters
pl1 | f1 channel – power level for pulse (default)
pl2 | f2 channel – power level for pulse (default)
pl9 | f1 channel – power level for presaturation
pl10 | f1 channel – power level for TOCSY-spinlock
pl11 | f1 channel – power level for ROESY-spinlock
pl12 | f2 channel – power level for CPD/BB decoupling
pl14 | f2 channel – power level for cw saturation
pl15 | f2 channel – power level for TOCSY-spinlock
sp1 | f1 channel – shaped pulse for selective excitation or f1 channel - shaped pulse for water flipback
sp2 | f1 channel – shaped pulse 180 degree or f2 channel - shaped pulse 90 degree (on resonance)
sp7 | f2 channel – shaped pulse 180 degree (off resonance2) or f2 channel – shaped pulse 180 degree (adiabatic) or f1 channel - shaped pulse for wet
p0 | for different applications i.e. f1 channel - variable flip angle high power pulse in DEPT
p1 | f1 channel - 90 degree high power pulse
p2 | f1 channel – 180 degree high power pulse
p3 | f2 channel - 90 degree high power pulse
p4 | f2 channel – 180 degree high power pulse
p6 | f1 channel - 90 degree low power pulse
p11 | f1 channel - 90 degree shaped pulse (selective excitation or water flipback/watergate or wet)
p15 | f1 channel – pulse for ROESY spinlock
p16 | homospoil/gradient pulse
p17 | f1 channel – trim pulse at pl10 or pl15
p18 | f1 channel – shaped pulse (off resonance presaturation)
d0 | incremented delay (2D) [3 usec]
d1 | relaxation delay 1-5 * T1
d2 | 1/(2J)
d3 | 1/(3J)
d4 | 1/(4J)
d6 | delay for evolution of long range couplings
d7 | delay for inversion recovery
d8 | NOESY mixing time
d9 | TOCSY mixing time
d11 | delay for disk I/O [30 msec]
d12 | delay for power switching [20 usec]
d14 | delay for evolution after shaped pulse

## Slide 15
Avance 1D/2D
BRUKER
23
d16 | delay for homospoil/gradient recovery
d17 | delay for DANTE pulse-train
d18 | delay for evolution of long range couplings
d19 | delay for binomial water suppression
d20 | for different applications
cnst0 | for different applications
cnst1 | J (HH)
cnst2 | J (XH)
cnst3 | J (XX)
cnst4 | J (YH)
cnst5 | J (XY)
cnst11 | for multiplicity selection
cnst12 | for multiplicity selection
vc | variable loop counter, taken from vc-list
vd | variable delay, taken from vd-list
l1 | loop for MLEV cycle (((p6*64) + p5) * l1) + (p17*2) = mixing time
l2 | loop for GARP cycle l2 * 31.75 * 4 * p9 => AQ
l3 | loop for phase sensitive 2D or 3D using States et al. or States-TPPI method l3 = td1/2
l4 | for different applications i.e. noediff
Note that the default units for pulses are microseconds (u), the units for delays are seconds (s), but one can always enter a value combined with a unit to define a time slot in XWinNMR. The nomenclature here is: s = seconds, m = milliseconds and u = microseconds. For example To set the value of d1 to 500m would define d1 to last for half a second.
The complete information on the nomenclature and default usage of the pulse program parameters can be found in:
$XWinNMRHome/exp/stan/nmr/lists/pp/Param.info
The nomenclature and description of the standard pulse programs and predefined parameter sets can be found in:
$XWinNMRHome/exp/stan/nmr/lists/pp/Pulprog.info
Acquisition, processing and plotting commands can be given either in the XWinNMR command line or via menu selection. Examples are zg, which starts the acquisition, ft which performs a Fourier transformation on the current data or apk which invokes the automatic phase correction.
Another possibility to manage different task in XWinNMR are AU programs. They handle many routine jobs an can be selected or edited by the edau command. AU programs have to be compiled before first usage. Compile and start AU Programs by entering xau followed by the program name.
XWinNMR also offers extensive online documentation, which can be accessed via the help menu in the XWinNMR menu bar.

## Slide 16
24
BRUKER
Avance 1D/2D
2.2.3 Changes for XWinNMR 3.5
XWinNMR version 3.5 is shipped with new systems now. There are some new commands and the handling of some pulse programs have changed from the software version 3.1.
In XWinNMR 3.5 the names of pulse program and parameter files have been adjusted to the general NMR nomenclature. For recording HSQC, HMQC and HMBC spectra pulse program and parameter files starting with the 4 letter code hsqc, hmqc, and hmbc, respectively, have to be given in the pulprog line in the eda table.
A new parameter TD0 is now available in the eda table. This parameter brings about a storage of your 1D data after recording ns/TD0 scans. This is especially useful for very long 1D experiments.
For more information on general changes, please refer to the release letter of your software packet. Information for pulse program specific changes can be found in: $XWinNMRHome/exp/stan/nmr/lists/pp/Update.info
2.3 Tuning and Matching the Probe
In a probehead there are resonant circuits for each nucleus indicated on the probehead label (e.g., one for 1H and one for 13C in a dual 1H/13C probehead; one for 1H and one for a wide range of nuclei in BBO or BBI probeheads). There is also a resonant circuit for the lock nucleus, but the standard user will never need to adjust this, so we will ignore it in the following. Each of the circuits has a frequency at which it is most sensitive (the resonance frequency). Once the sample is inserted, the probehead should be tuned and matched for these individual frequencies.
Tuning is the process of adjusting this frequency until it coincides with the frequency of the pulses transmitted to the circuit. For example, the frequency at which the 1H resonant circuit is most sensitive must be set to the carrier frequency of the 1H pulses (which is sfo1 if the 1H circuit is connected to the f1 channel, sfo2 if it is connected to the f2 channel, etc.). Matching is the process of adjusting the impedance of the resonant circuit until it corresponds with the impedance of the transmission line connected to it. This impedance is 50 . Correct matching maximizes the power that is transmitted to the coil. A probehead is said to be tuned and matched when all of its resonant circuits are tuned and matched. Once a probehead has been tuned and matched, it is not necessary to retune or rematch it after slight adjustments of the carrier frequency, since the probehead is generally tuned and matched over a range of a couple of hundred kHz. On the other hand, large adjustments to the carrier frequency, necessary when changing nuclei, warrant retuning and rematching of the probehead. Thus, a broadband probe needs to be retuned and rematched each time the heteronucleus is changed.

## Slide 17
Avance 1D/2D
BRUKER
25
If you have an ATM probe, enter edsp and set the spectrometer parameters for the channels that should be matched and tuned. For 1H on channel F1 and 13C on channel F2 enter the following values:
NUC1 NUC2 NUC3
1H
13C
OFF
This automatically sets sfo1 to a frequency appropriate for 1H and sfo2 to the corresponding 13C frequency for tuning and matching. Exit edsp by clicking SAVE.
Type atma. This will invoke the automatic match and tune program for all nuclei that were selected previously in edsp. Therefore it is not necessary to tune and match manually.
2.4 Tuning and Matching 1H (non ATM Probes)
When the NMR experiments to be performed are
1H
homonuclear
experiments (e.g., 1H 1D spectroscopy, COSY, NOESY, or TOCSY), only the
1H circuit of the probehead has to be tuned and matched.
Make sure that the sample is in the magnet, and the probehead is connected for standard 1H acquisition. Note that there is no special configuration for tuning and matching. Also, it is recommended to tune and match without sample spinning.
2.4.1 Set the Parameters
In XWIN-NMR, enter edsp and set the following spectrometer parameters:
NUC1 NUC2 NUC3
1H OFF OFF .
This automatically sets sfo1 to a frequency appropriate for 1H tuning and matching. There is no need to adjust sfo1 carefully now. Exit edsp by clicking SAVE.
2.4.2 Start Wobbling
Tuning and matching are carried out simultaneously using XWIN-NMR. During wobbling, a low power signal is transmitted to the probehead. This signal is swept over a frequency range determined by the parameter wbsw (the default value is 4 MHz) centered around the carrier frequency (i.e., sfo1, sfo2, etc., depending on which nucleus is being tuned/matched). Within the preamplifier (High Performance Preamplifier Assembly or HPPR), the impedance of the probe over this frequency range is compared to the impedance of a 50  resistor. The results are shown both on the LED display of the HPPR and in the acquisition submenu of XWIN-NMR. Both displays show the reflected power of the probehead versus the frequency of the

## Slide 18
26
BRUKER
Avance 1D/2D
signal. The user observes either one or both of these displays while tuning and matching the probehead.
Before starting the wobbling procedure, ensure that no acquisition is in progress, e.g., enter stop.
Enter acqu to switch to the acquisition window of XWIN-NMR, if it is desired to use this to monitor the tuning and matching. Notice that being in the acquisition window slows down the wobbling procedure, so if the HPPR LED display will be used to monitor tuning and matching, it is best to remain in the main XWIN-NMR window and not to switch to the acquisition window.
Start the frequency sweep by typing wobb. The curve that appears in the acquisition window is the reflected power as a function of frequency. Unless the probehead is quite far from the correct tuning and matching, there will be a noticeable dip in the curve. When the 1H circuit is properly tuned, the dip will be in the center of the window, denoted by the vertical marker; and when the circuit is properly matched, the dip will extend all the way down to the x axis. Similar information is conveyed by the LED display on the HPPR. The horizontal row of LED's indicates tuning and the vertical row matching. When the circuit is properly tuned and matched, the number of LEDs is minimized. This usually means that only green LED, are lit in both the horizontal and vertical displays.
2.4.3 Tune and Match
Adjust the tuning and matching screws (labeled T and M) at the base of the probehead. Note that the screws are color coded and those for the 1H circuit are usually yellow. Also note that the screws have a limited range and attempting to turn them beyond this range will damage the probehead.
Since there is an interplay between tuning and matching, it is generally useful to adjust the T and M screws in an iterative fashion. Turn the M screw until the dip is well matched at some frequency (the dip extends to the x axis and the number of LEDs lit in the vertical HPPR display is minimized). Most likely this will not be the desired frequency. Adjust the T screw slightly to move the dip toward the center of the window, or equivalently, to reduce the number of LEDs lit in the horizontal HPPR display. Rematch the dip by adjusting the M screw again. Note that it is possible to run out of range on the M screw. If this happens, return M to the middle of its range, adjust T to get a well matched dip at some frequency, and walk the dip towards the correct frequency as described above.
As mentioned above, ideal tuning and matching is when the dip is centered in the window and extends to y = 0 (the x axis) on the acquisition window, or equivalently, when the number of LED's lit on the preamplifier is minimized in both the vertical and horizontal display.
When the 1H circuit is tuned and matched, exit the wobble routine by typing stop. Click on return to exit the acquisition window and return to the main window.

## Slide 19
Avance 1D/2D
BRUKER
27
  Tuning and Matching 13C (non ATM Probes)
Since most 13C experiments make use of 1H decoupling, besides 13C the 1H should be tuned and matched as well. When tuning and matching a probehead with multiple resonant circuits, it is best to tune and match the lowest frequency circuit first. Thus, when tuning and matching a probehead for both 1H and 13C, first do the 13C and then the 1H adjustments.
Make sure that the sample is in the magnet, and the probehead is connected for the appropriate experiment. Also, it is recommended to tune and match without sample spinning.
    Set the Parameters
In XWIN-NMR, enter edsp and set the following spectrometer parameters:
NUC1 NUC2 NUC3
13C OFF OFF .
This automatically sets sfo1 to a frequency appropriate for 13C tuning and matching. Exit edsp by clicking SAVE.
2.5.2 Start Wobbling, Tune and Match
Ensure that no acquisition is in progress, enter stop.
Enter acqu to switch to the acquisition window, if this will be used to monitor the tuning and matching.
Start the frequency sweep by typing wobb. The curve that appears in the acquisition window is for 13C. Adjust the tuning and matching following the guidelines given above for 1H. Notice that some probeheads (e.g., broadband probeheads) have sliding bars instead of screws, one set labeled tuning and another labeled matching. Set the tuning and matching sliding bars to the values indicated for 13C on the menu. Adjust the tuning and matching bars until the dip is well tuned and matched at some frequency as described above for 1H.
Once the 13C circuit is tuned and matched, the 13C wobbling must be stopped before the 1H wobbling. Exit the wobble routine by typing stop. Enter edsp, change NUC1 to 1H, and exit by clicking SAVE. Start the 1H frequency sweep by typing wobb. After a few seconds the 1H curve appears in the acquisition window and the 1H circuit can be tuned and matched as described above.
Alternatively, if the user already has a data set in which NUC1 = 1H and NUC2 = OFF, there is no need to redo edsp for the current data set. The user may simply read in the 1H data set and then type wobb.
Once the probehead is tuned and matched for 13C and 1H, exit the wobble routine by typing stop.
Click on to exit the acquisition window and return to the main window.

## Slide 20
28
BRUKER
Avance 1D/2D
  Locking and Shimming
Before running an NMR experiment, it is necessary to lock and shim the magnetic field.
    Locking
To display the lock signal enter lockdisp. This opens a window in which the lock trace appears.
The most convenient way to lock is to use the XWIN-NMR command lock. To start the lock-in procedure, enter lock and select the appropriate solvent from the menu. Alternatively, enter the solvent name together with the lock command, e.g., lock cdcl3. During lock-in, several parameters such as the lock power, the field value, and the frequency shift for the solvent are set according to the values in the lock table. This table can be edited using the command edlock. Note that the lock power listed in this table is the level used once lock-in has been achieved. The field-shift mode is then selected and autolock is activated. Once lock-in is achieved, the lock gain is set so that the lock signal is visible in the lock window. At this point the message “lock: finished” appears in the status line at the bottom of the window.
The lock-in procedure outlined above sets the frequency shift to the exact frequency shift value for the given solvent as listed in the edlock table. It also sets the field value to the value listed in the edlock table and then adjusts it slightly to achieve lock-in (the absolute frequency corresponding to a given ppm value no longer depends on the lock solvent). Following this lock-in procedure, the solvent parameter in the eda table is set automatically, which is important if you wish to use the automatic calibration command sref (see “Spectrum Calibration and Optimization”).
The lock-phase adjustment by monitoring the sweep wiggles (i.e., while the field is not locked but is being swept) is recommended each time the probehead is changed, because autolock may fail. If the original phase is reasonably close to the correct value, lock-in can be achieved and the phase can be adjusted using autophase. Note that the lock phase for each probehead is stored in the edlock table. In some cases, the lock power level listed in the edlock table is set too high leading to a saturation of the lock signal. Usually, lock-in can be achieved, but the signal oscillates due to saturation. A quick fix is simply to reduce the lock power manually after lock- in. However, it is better to change the power level in the edlock table. Note that the appropriate lock power level depends on the lock solvent, the field value, and the probehead.
    Shimming
If the sample has been changed, the first step after locking is shimming the magnetic field. Enter rsh and select an appropriate shim file from the menu. Usually, only the Z and Z2 shims (and probably the X and Y) must be adjusted while observing the lock signal. The best shim values correspond to the highest lock level (height of the lock signal in the window). For further

## Slide 21
Avance 1D/2D
BRUKER
29
discussion of shimming see Chapter 6 ‘Shim Operation' of the BSMS User's Manual.
If you have a gradient probe, you can also use the gradient shimming tool, which can be started by the command gradshim. For more Information, please refer to the gradient shimming installation and users guide which is available online in the XWinNMR help menu.
2.6.3 Optimize lock settings (optional)
Once the magnetic field has been locked and shimmed, the user may wish to optimize the lock settings as described below. It is strongly recommended to follow this procedure before running any experiment requiring optimal stability (e.g., NOE difference experiments).
After the field is locked and shimmed, start the auto-power routine from the BSMS keyboard (see Chapter 2 ‘Key Description' of the BSMS User's Manual). For lock solvents with long T1 relaxation times (e.g., CDCl3), however, auto power may take an unacceptably long time and the lock power should be optimized manually. Simply increase the lock power level until the signal begins to oscillate (i.e., until saturation), and then reduce the power level slightly (approximately 3 dB). For example, if the lock signal begins to oscillate at a power of –15 dB, the optimal magnetic field stability can be expected when a level of approximately –18 dB (or even –20 dB) is used. The field stability will be significantly worse if a power level of, say, –35 dB is used instead.
When the lock power is optimized, start the auto-phase routine, and finally the auto-gain routine. Take note of the gain value determined by auto gain. Using this value, select the appropriate values for the loop filter, loop gain, and loop time as shown below in Table 12.

## Slide 22
30
BRUKER
Avance 1D/2D
Table 12: Lock Parameters (BSMS Firmware Version 980930)
Lock RX Gain (after auto gain) [dB] | Loop Filter [Hz] | Loop Gain [dB] | Loop Time [sec]
120 | 20 | –17.9 | 0.681
 | 30 | –14.3 | 0.589
110 | 50 | –9.4 | 0.464
 | 70 | –6.6 | 0.384
 | 100 | –3.7 | 0.306
 | 160 | 0.3 | 0.220
 | 250 | 3.9 | 0.158
 | 400 | 7.1 | 0.111
90 | 600 | 9.9 | 0.083
 | 1000 | 13.2 | 0.059
 | 1500 | 15.2 | 0.047
 | 2000 | 16.8 | 0.041
So, for example if auto gain determines a lock gain of 100 dB, the user should set the loop filter to 160 Hz, the loop gain to 0.3 dB, and the loop time to 0.220 sec (see Chapter 4 ‘Menu Description' of the BSMS User's Manual for how to set these parameters from the BSMS keyboard).

## Slide 23: 3 Basic 1H Acquisition and Processing
Avance 1D/2D
BRUKER
31
3.1 Introduction
This chapter describes the acquisition and processing of a 1D 1H NMR spectrum using the simple one-pulse NMR experiment shown in Figure 1. The pulse sequence consists of the recycling delay, trd, the radio-frequency (RF) pulse, and the acquisition time during which the signal is recorded. The pulse angle is shown to be /2, although, in practice, it is often chosen less. The two parameters, d1 and p1, correspond to the length of the recycle delay, trd, and the length of the RF pulse, respectively.
Note that the time intervals depicted in the pulse sequence diagrams are not drawn to scale. For example, d1 is typically a few seconds while p1 is typically a few microseconds in length.
Figure 1: 1D 1H NMR One-Pulse Sequence
/2
¦
d1
¦ p1
acq
1H
trd
    Sample
The sample used for demonstrating the basic 1D 1H experiment is 100 mg Cholesterylacetate in CDCl3 with 0.5 % TMS.
    Preparation
Make sure that you have done the following steps (see also Chapter 2 ‘Preparing for Acquisition'):
      Insert a suitable probehead
      Read in the corresponding shim file
      Insert the sample
      Lock the spectrometer

## Slide 24
32
BRUKER
Avance 1D/2D
Optimize the Z and Z2 (and probably X and Y) shims
Tune and match the probehead for 1H
3.2 Spectrometer and Acquisition Parameters
Before the acquisition of a spectrum a new data set must be created. All the spectrometer and acquisition parameters are entered within the new data set.
The spectrometer parameters are responsible for the hardware settings necessary for configuring the spectrometer for a particular experiment. The command edsp calls up a window in which the spectrometer parameters for
the observe and the decoupler channel(s) are set.
The acquisition parameters include all pulse sequence parameters, the number of data points, number of scans, receiver gain, and many others. These may be displayed and edited by entering eda. Notice that the spectrometer parameters are also listed in the eda table. It is important to set the spectrometer parameters before setting the acquisition parameters, because the values from edsp automatically overwrite the corresponding ones from the eda table.
3.3 Create a New File Directory for the Data Set
To create a new data set, type edc in the command line of the XWIN-NMR window. This calls up a small window entitled “Current Data Parameters”. Enter a data set name (NAME), an experiment number (EXPNO), a processed data number (PROCNO), the disk unit (DU) where the data is stored, the user id (USER), and the data type (TYPE). Change the parameters as follows:
NAME EXPNO PROCNO
proton 1
1
Click on SAVE. This exits edc and creates the data set proton/1/1. The message “NEW DATA SET” should now appear on the screen.
3.4 Set Up the Spectrometer Parameters
Enter edsp and set the following spectrometer parameters:
NUC1 NUC2 NUC3
1H
off off

## Slide 25
Avance 1D/2D
BRUKER
33
Since there is no decoupling, the only relevant spectrometer parameters are SFO1. Click on SAVE to save the spectrometer parameters and return to the main window. The spectrometer is now prepared to pulse and detect at the 1H frequency.
3.5 Set Up the Acquisition Parameters
Enter eda and set the acquisition parameters as shown in Table 13, where only the relevant parameters are listed. Note that the parameters d1, p1, and pl1 are included in the parameter arrays D, P and PL in the eda table, respectively. These parameters can be edited within eda, by clicking the ‘
**Array**-button next to the corresponding parameter
Table 13: Basic 1H Spectrum Acquisition Parameters
Parameter | Value | Comments
PULPROG | zg | see Figure 1 for the pulse sequence diagram
AQ_mod | DQD | If DQD is not available, use qsim
TD | 32 k | 32 k is a standard value for a high-resolution 1D spectrum
PARMODE | 1D | One-dimensional experiment
NS | 1 | one scan is recorded for parameter optimization
DS | 0 | no dummy scans are recorded
D1 | 2 | the default unit for delays is seconds; entering “2” sets a delay of 2 seconds (click the D**Array** button)
P1 | 3 | the default unit for pulse lengths is microseconds; entering “3* sets a pulse length of 3 microseconds (s) (click the P**Array** button)
PL1 | PL1 = | power level for the p1 pulse see also “An Important Note on Power Levels” on page 3 (click the P**Array** button)
SW | 50 | for the first spectrum of an unknown sample use a large spectral width; when you enter “50” the registered value is slightly different
RG | 64 | suggested receiver gain
NUC1 | 1H | observe nucleus
O1P | 15 | position of the carrier frequency is 15 ppm

## Slide 26
34
BRUKER
Avance 1D/2D
Click on SAVE to save the acquisition parameters and return to the main window. Click on DONE to save the changes and return to the eda table. As with most acquisition parameters, however, d1, p1, and pl1 can also be edited by typing them in the command line of the main XWIN-NMR window. As mentioned before, most of the acquisition parameters for the current pulse program can also be entered in the ased table.
3.6 Acquisition
Enter acqu to switch to the acquisition window. While it is possible to acquire a spectrum from the main window, the buildup of the FID can only be observed in the acquisition window.
Enter the command rga, which performs several acquisitions and sets a suitable value for the receiver gain (rg). Enter zg, which deletes any previous data (‘zero') and starts the experiment (‘go'). The message scan 1/1 indicates that the spectrometer is performing the first scan and that only one scan will be performed.
If, at any time, a submenu is entered accidentally, click on the button located on the button bar and then enter acqu to switch back to the acquisition window.
If, at some point the message “DATA OUT OF WINDOW ” appears, or if the scaling is unsuitably large or small, click on the and buttons located on the button bar.
3.7 Processing
After the FID has been acquired the next step is to process the acquired data. The processing parameters are displayed and edited by entering edp. First, Fourier transformation is performed by entering the command ft. The number of points used to resolve the resulting spectrum is determined by the processing parameter si (size). The spectrum consists of si real points and si imaginary points, and thus the default setting of si is td/2, where td is the acquisition parameter indicating the number of time domain data points. In general, td/2 and si are numbers described by powers of 2 (2, 4, 8, 16, 32, 64, 128, ...). If si < td/2 not all the time domain data is used for the Fourier transformation, and if si > td/2 the time domain data is zero-filled with 2(si) before the Fourier transformation. In 1D spectroscopy, it is often recommended to zero-fill once, i.e., to set si = td.
Check the value of si. Enter si and when prompted enter 32k (appropriate for td = 32 k). Enter ft: The display automatically switches from the
acquisition window to the main window and displays the. The FID can still be viewed by returning to the acquisition window. If the x axis of the Fourier transformed spectrum is displayed in Hz, click on to convert into a ppm scale. If necessary, use the buttons as described above to scale the spectrum.

## Slide 27
Avance 1D/2D
BRUKER
35
You can zoom into a part of the spectrum by defining the appropriate 1D plot range. Move the cursor into the display window and press the left mouse button to tie the cursor to the spectrum. Move the cursor to one side of the desired zoom region and click the middle mouse button to define it. Move the cursor to the other side of the desired plot region and click the middle mouse button again to zoom into this region. To display the whole spectrum push the button.
3.8 Phase Correction
Once the spectrum is Fourier transformed it must be phase corrected. Click on to enter the phase correction submenu. Click on for setting the reference for the 0th-order phase correction to the position of the biggest peak in the spectrum and adjusts its phase. To adjust the 0th-order phase manually, place the cursor on and hold down the left mouse button. Move the mouse until the reference peak is positive and the baseline on either side is as flat as possible.
Most likely, the peaks on either side of the reference peak are not yet phased correctly and require a 1st-order phase correction. To adjust the 1st-order phase correction, place the cursor on and hold down the left mouse button, and move the mouse until the peaks far from the reference point are also in-phase.
Note that it is advisable to select the reference peak for the 0th-order phase correction near one edge of the spectrum. However, for some samples the biggest peak will be located in the middle of the spectrum. In such cases,
click on and define the reference peak by moving the cursor onto
the desired peak and clicking with the middle mouse button.
Once the spectrum is phased correctly, click on return to exit the submenu and save the phase corrections by selecting Save & return. The 0th- and 1st- order phase correction values are stored as processing parameters phc0 and phc1, respectively. To quit the phase correction submenu without saving the corrections, simply click on return and select return. In either case, the display returns to the main menu and the spectrum appears on the screen.
Note that once suitable values of phc0 and phc1 have been stored it is possible to use them for phase correcting subsequent spectra by typing the command pk. In addition, the Fourier transformation (ft) and the phase correction (pk) can be performed within one step using the command fp.
3.9 Windowing
Before the Fourier transformation is performed, it is common to apply a window (or filter) function to the time domain data. The main reason for this is the improvement of either signal-to-noise, or resolution. Usually, for a simple 1D spectrum as described here, the signal-to-noise ratio is improved by multiplying the FID with a simple exponential function achieved by the command em.

## Slide 28
36
BRUKER
Avance 1D/2D
The decay rate of the exponential function determines the amount of line broadening. This rate is determined by the processing parameter lb (in Hz). Enter lb and set the value to 0.3, which corresponds to an appropriate line
1H
broadening for high-resolution spectra. Enter em to perform the
multiplication, and then enter fp to Fourier transform and phase correct the filtered data. You can also use the combined command efp, which performs the windowing, Fourier transformation and the phasing with the previously determined phase correction. The final spectrum should look like the one shown in Figure 2.
Figure 2: 1H 1D Spectrum of 100 mg Cholesterylacetate in CDCl3
9.0
8.0
7.0
6.0
5.0
4.0
3.0
2.0
1.0
0.0
ppm

## Slide 29
Avance 1D/2D
BRUKER
37
3.10 Integration
To quantitatively analyze an observed signal, the integrated intensity of the peaks are compared within each other.
Click to enter the integration submenu. To integrate a peak, first move the cursor into the spectral window and click the left mouse button. Next, click the middle mouse button once at each end of the range of interest; the integral appears automatically. Click the left mouse button again to release the cursor from the spectrum. An asterisk or a vertical arrow appears next to the right end of the integral (if not, select the integral with the left mouse button). Correct the baseline of the integral with the
buttons. Integrate the other areas or peaks in the same way.
For the calibration, select an integral (asterisk/arrow) and click on
Enter 100 to calibrate this integral to 100%. Upon return select Save & store ‘intrng' to save the integral and normalization constant and return to the
main 1D processing window.
It is also possible to compare integral values of spectra located in different data sets: Integrate both spectra and calibrate the integral(s) in one of them,
e.g. to 100 as described above. Enter the integration mode in the second spectrum, select the corresponding integral (asterisk/arrow) and click on the button to display the integral value compared to the calibrated
100% of the other signal.

## Slide 30
38
BRUKER
Avance 1D/2D

## Slide 31: 4 Pulse Calibration: Protons
Avance 1D/2D
BRUKER
39
4.1 Introduction
This chapter describes pulse calibration procedures for 1H and 13C. It is assumed that the user is already familiar with acquisition and processing of simple 1D NMR spectra. Appendix A (Data Sets and Selected Parameters), which lists all data sets generated throughout this course, and Appendix B (Pulse Calibration Results), which provides all the pulse lengths and power levels determined during this course, maybe useful in this context.
4.2 Proton Observe 90° Pulse
For the calibration of a 1H 90° pulse on the observe channel (F1), the one- pulse sequence described in Chapter 3 is used. The carrier frequency (o1p) is set onto the resonance frequency of a peak in the 1H spectrum of an appropriate sample. This peak is monitored while the length (p1) and/or the strength (pl1) of the RF pulse is adjusted to determine the exact conditions for a 90° pulse.
A common sample used for the 1H pulse calibration is 0.1% Ethylbenzene in
1H
CDCl3. Ethylbenzene shows a simple spectrum with well-separated
signals, which facilitates the selection of a single resonance line. However, due to the relatively long spin-lattice or longitudinal relaxation time (T1) of Ethylbenzene, a long recycle delay time must be used.
4.2.1 Preparation
Insert the sample and lock the spectrometer (lock). Readjust the Z and Z2 shims until the lock level is maximal (use lockdisp). Tune and match the probehead for 1H observation (see Chapter 2.3).
First, create a new data set. Since this will be a 1H observe experiment, it is best to start out from a previous 1H data set, e.g., proton/1/1: Enter re proton 1 1, then enter edc and change the following parameters:
NAME test1h
EXPNO 1
PROCNO 1
Click on SAVE to create the data set test1h/1/1.
Enter eda and set the acquisition parameter values as shown in Table 14.

## Slide 32
40
BRUKER
Avance 1D/2D
Table 14: 1D 1H one-pulse Acquisition Parameters
Parameter | Value | Comments
PULPROG | zg | see Figure 1 for the pulse sequence diagram.
TD | 4 k | 
NS | 1 | number of scans
DS | 0 | no dummy scans
D1 | 10 | interscan delay (10s, because of long T1)
P1 | 3 | start with 3s, which should correspond to less than a 90° pulse
PL1 |  | power level for the p1 pulse see “An Important Note on Power Levels” on page 3
SW | 20 | start with a large spectral width of 20ppm; which will be optimized lateron
o1p | 5 | will be optimized lateron
Enter rga to perform an automatic receiver gain adjustment, then enter zg to acquire the FID, and edp to set the processing parameters as shown in Table 15.
Table 15: 1D 1H one-pulse Processing Parameters
Parameter | Value | Comments
SI | 2 k | 
LB | 1 Hz | 
PSCAL | global | 
Fourier transform the spectrum with the command ef and phase the spectrum according to Chapter 3.8. Type sref to calibrate the spectrum and confirm the message “no peak found in ‘sref' default calibration done”.
4.2.2 Optimize the Carrier Frequency and the Spectral Width
The carrier position (o1p) should now be set to the signal used for monitoring the 90° pulse calibration, which is the quartet signal of the Ethylbenzene 1H spectrum. Expand the spectrum so that only the quartet at 2.6 ppm is displayed. Click on to enter the calibration submenu. Click on with the left mouse button, move the cursor to the center of the quartet and click the middle mouse button to assign o1p to this frequency. Click on
to exit the calibration submenu and return to the main window. Reduce the spectral width by entering swh and change the value to 1000 (Hz).

## Slide 33
Avance 1D/2D
BRUKER
41
Enter zg to acquire a new FID using the new values for o1p and swh and process the spectrum using the command ef.
4.2.3 Define the Phase Correction and the Plot Region
The phase correction and the spectral region plotted in the output file must be optimized before the automation program for the pulse calibration is executed. Phase correct the spectrum according to Chapter 3.8 in a way that the quartet signal is positive. Expand the spectrum so that the quartet covers approximately the central quarter of the screen. Click on with the left mouse button and hit return for the following three prompts, or answer them as follows:
F1 F2
change y-scaling on display according to PSCAL?
2.8 ppm
2.4 ppm y
The preparations are now completed and the calibration experiment can be executed as described in the next section.
4.2.4 Calibration: High Power
For the 90° pulse calibration an automation program called paropt is used. (Since the execution of this automation is time consuming, it is not the best choice if the correct pulse times and power levels are already known approximately. In such cases, the correct values are usually just checked by acquiring 1D spectra with different pulse widths to check for maximal signal.)
The automation program is started by typing xau paropt and answering the appearing questions as follows:
Enter parameter to modify: Enter initial parameter value: Enter parameter increment: Enter # of experiments:
p1 2
2
16
The spectrometer acquires and processes 16 spectra with incrementing the parameter p1 from 2 sec by 2 sec to a final value of 32 sec. For each of the 16 spectra, only the spectral region defined above is plotted, and all the spectra are plotted side-by-side in the file test1h/1/999 as shown in Figure 3. At the end of the experiment, the message “paropt finished” and a value for the parameter p1 is displayed, which corresponds to the 90° pulse length of the 1H transmitter with the power level as defined by pl1. Write this value down and follow the procedure described below to obtain a more accurate 90° pulse measurement.
Return to the data set test1h/1/1 by entering re 1 1. Type p1 and enter a value which corresponds to a 360° pulse (i.e., four times the 90° value determined by paropt before). Acquire and process a new spectrum by typing zg and efp (see Chapter 3.9) respectively. Change p1 slightly and acquire and process a spectrum again, until the quartet undergoes a zero crossing as expected for an exact 360° pulse. Note that the quartet signal is negative for pulse angles slightly less than 360° and positive when the pulse angle is slightly more than 360°.

## Slide 34
42
BRUKER
Avance 1D/2D
The 360° pulse length divided by four yields the accurate 1H 90° transmitter pulse length for the actual power level pl1.
Figure 3: Paropt Results for 1H 90° Pulse Calibration
p1 = 180°
p1 = 90°
p1 = 270°
4.2.5 Calibration: Low Power for MLEV Pulse Train (TOCSY)
The 1H 90° pulse for the MLEV pulse train used during the spinlock period of a TOCSY sequence is between 30 sec to 40 sec. The procedure outlined below uses the paropt routine to determine the corresponding power level. However, the power level can be estimated roughly by using a rule of thumb: The pulse length doubles for an additional 6 dB increase of the power level. For example, the 90° pulse length (p1) was determined 8 sec for pl1 = 0 dB. Thus, the p1 = 16 sec for pl1 = 6 dB, or the p1 = 32 sec for pl1 = 12 dB.

## Slide 35
Avance 1D/2D
BRUKER
43
For performing the exact determination of the low power pulse, return to the file test1h/1/1 (re 1 1). Enter p1 and change the value to 35 (sec), type xau paropt and answer the questions as follows:
Enter parameter to modify: Enter initial parameter value: Enter parameter increment: Enter # of experiments:
pl1 0
1
16 .
Again, the 16 spectra will be displayed in the file test1h/1/999 and at the end of the experiment, the message “paropt finished” and a value for pl1 is displayed. This value corresponds to the 1H transmitter power level for a 90° pulse length of 35 sec. Write down this value and follow the procedure described below to obtain a more accurate 90° pulse measurement.
Return to test1h/1/1 (re 1 1), type p1 and change the value to 140 sec (= 360° pulse). Acquire and process a spectrum (zg, efp) by using the power level pl1 determined by paropt above. Change pl1 slightly until the quartet undergoes a zero-crossing indicating the accurate 360° pulse. Divide this 360° pulse time by four to get the 90° pulse length.
Note that the parameters used by the TOCSY sequence are p6 for the 90° pulse length and pl10 for the power level, rather than p1 and pl1.
4.2.6 Calibration: Low Power for ROESY Spinlock
The power level required for the cw spinlock pulse used with ROESY experiments corresponds to a 90° pulse length of 100 sec to 120 sec. As described for the 90° pulse determination of the MLEV pulse above in Chapter 4.2.5, the power level can again be estimated using the rule of thumb, or measured using the paropt automation.
When using paropt, return to the file test1h/1/1 (re 1 1), enter p1 and change the value to 110 (sec), and type xau paropt. Answer the questions as follows:
Enter parameter to modify: Enter initial parameter value: Enter parameter increment: Enter # of experiments:
pl1 10
1
16 .
The results are displayed in the file test1h/1/999, and at the end of the experiment, the message “paropt finished” and a value for pl1 corresponding to the 1H transmitter power level for a 90° pulse length of 110
sec are displayed. Follow the same procedure as described in Chapters
4.2.4 and 4.2.5 for a more accurate determination of the power level.
Note that since ROESY uses cw spinlock, only the power level determination is important here, but not the actual 90° pulse length. The power level parameter used with the ROESY sequence is pl11, rather than pl1.