## Slide 1: Water Suppression Techniques
[H2O]=55,000 mM [Protein]< 5 mM
[H2O]/[Protein]>11,000
Water suppression
Sample used throughout this lecture: 1 mM TEP-I in 90% H2O/10%D2O, pH 6.0, 290 K.

## Slide 2: Water Suppression Technique
Presaturation
Watergate
Water flip-back
Jump and return, 1-1, 1331
Suppression by coherence pathway rejection

## Slide 3: Water Suppression Technique in Protein NMR
Labile, exchange with water
(pH, structure, temperature dependent)
Resonates near
the water frequency
C
O
O
N
H
C
H
C
C
Backbone NH
aromatic
Side-chain NH2
Ha
aliphatic
1H

## Slide 4: pH Dependence of Amide Proton Exchange Rates
~330 times faster
In practice: the pH value for a protein sample for NMR studies is kept below 7.5 to avoid fast exchange rates.
Figure modified from p154 of John Cavanagh et al.,“Protein NMR Spectroscopy: Principles and Practice”, Academic Press (1995)
C
O
N
H
C
C
H
C
O
N-H + *H2O N-*H + *H-O-H

## Slide 5: Presaturation
90x
Acqu
cw
d1
Apply a low power C.W. irradiation
on water before the first 90 degree pulse, usually during the relaxation delay
Parameter adjustment: Pulprog=zgpr
Adjustment: pl9; power level for presaturation
1D 1H spectra of TEP-I in 90%H2O/10%D2O
, pH 6, 290 K.

## Slide 6: Drawback of Presaturation
Saturation transfer to exchangeable NH protons
Bleaching of signals near water
Large dispersive tail of water signal: tilted baseline
Watergate (zgpgwg)
Presat (zgpr)
Watergate
Presat
1D 1H spectra of TEP-I, pH 6, 290 K.

## Slide 7: Pulsed Field Gradient (PFG)
A field-gradient pulse is a pulse or a period during which the magnetic field is made deliberately inhomogeneous.
B=B0+Bg(z)
The magnetic field, generated by a gradient pulse, Bg(z) varies linearly along the Z-axis
Bg(z)=zGz, where
Gz: gradient strength (G/cm), Z: z-axis position
Viewing on the rotating frame, spins at different z-position acquire different phase (Larmor frequencies): (z)=zGz where =phase, : gyromagnetic ratio, : gradient duration
Actively shielded gradient coil reduces eddy current, and is now popular in multidimensional NMR spectroscopy.
Z=0
Bg
-Bg

## Slide 8: Pulsed Field Gradient (PFG)
Z=0
Bg
-Bg
Gz
-Gz
Initially spins in each slice (isochromat) are “phase-coherence”.
After a field-gradient pulse, the spins at different slice experience different magnetic field strength, and acquire different Larmor frequencies. The “phase-coherence between slices is now lost due to Larmor precession.
The coherence can be refocused by another gradient pulse (gradient echo).
B
A C

## Slide 9: Pulsed Field Gradient (PFG)
Z=0
Bg
-Bg
Viewing from the Z-axis:
Gz
-Gz
B
A C
(figure from p106 of Sattler et al. Prog. In Nucl. Mag. Reson. Spect. 34 (1999)

## Slide 10: WATERGATE
90x
180y
d1
90-y
90-y
1H
Gz
a
b
Acqu

## Slide 11: WATERGATE
A strong Z-gradient pulse can be used to destroy transverse magnetization.
A destroyed (dephased) magnetization can be refocused by another z-gradient pulse of the same amplitude but of opposite phase. (or use a 180 pulse in between the two identical z- gradient pulse).
H2O: the two extra selective 90 pulse on water makes the 2nd z- gradient pulse act as another defocus gradient pulse.
Protein signals: the 180 pulse makes the 2nd Z-gradient act as a refocus gradient.
Ref: M. Piotto, V. Saudek & V. Sklenar, J. Biomol. NMR 2, 661 - 666 (1992)
90x 180y
d1
1H
Gz
a
90-y 90-y
b
Acqu

## Slide 12: WATERGATE
Parameter adjustment, Pulprog=zgpgwg
p11: pulse length for 90 degree shaped pulse
sp1: power level for 90 degree shaped pulse
spnam1: name of shaped pulse
For example: set spnam1=Sinc1.1000, p11=1 msec, Minimize the water fid by adjusting sp1 in the “gs” utility.
90x 180y
d1
90-y
90-y
1H
Gz
a
b
Acqu

## Slide 13: WATERGATE V.S. Presaturation
Watergate (pulprog=zgpgwg)
Presat (pulprog=zgpr)
1D 1H s
Wat
pectra of T
ergate
EP-I, pH 6,
Presat
290 K.
Watergate (zgpgwg)

## Slide 14: 3-9-19 WATERGATE
90x
d1
1H
Gz
Acqu
Sklenar et al., J. Magn. Reson., A102, 241-245 (1993)
Off resonance DANTE excitation technique.
3-9-19: 3--9--19-19--3 where 26=180, =delay.
(This is also referred as “W3”.)
3-9-19
500 MHz
Delay  =1/(4 max),
where 2max=distance of next null (Hz). (The delay  is field-dependent !!)

## Slide 15: 3-9-19 WATERGATE
Delay  =1/(4 max), where 2 max= distance of next null (Hz). (The delay  is field-dependent !!)
For example: Have the center of NH region (i.e. 8.2 ppm) to be the center of maximal excitation region:
=1/[4*(8.2-4.75)*600.13]=121 usec @600 MHz machine
=1/[4*(8.2-4.75)*500.13]=145 usec @500 MHz machine
Parameter adjustment: Pulprog=p3919
Set pl18=pl1, p27=p1, p0=p1
;d19: delay for binomial water suppression
;d19 = (1/(2*d)), d = distance of next null (in Hz)
Adjust d19 according to the magnetic field strength and where you want the center of maxima excitation to be.

## Slide 16: Water Flip-back WATERGATE
Parameter adjustment:
Pulprog=*fp*, i.e. “hsqcetfpf3gp” calibrate the shaped pulse as describe in WATERGATE.
d1
90-y
Gz
1H
a
b
Acqu
90x 180y
90X 90-y
S. Grzesiek and A. Bax, J. Am. Chem. Soc., 115, 12593-12594 (1993)
Water is aligned along the z axis before any z-gradient pulse (point “a”). So, it is not destroyed by the z-gradient pulse.
This reduces the signal loss of exchangeable protons due to attenuation of water signal (saturation transfer).

## Slide 17: W4, W5 WATERGATE
Better frequency response at the expense of pulse sequence length
(More Exchange)
Set  (d19) same as for 3-9-19
Simulated excitation profiles
Phases in the 1st and 2nd half differ by 180

## Slide 18
Water signal gets saturated
Excitation Sculpting

## Slide 19: Pulse Sequence for Observing Fast- Exchanging Protons
Flip-back WATERGATE (marginal performance)
Jump and return 1-1
1-3-3-1
N-H + *H2O N-*H + *H-O-H
Imino protons in DNA, hydroxyl protons (-OH), Histidine side chain protons in proteins are usually in a fast exchange process with water.

## Slide 20: Jump and Return: 1-1
Water signal: “on resonance”, aligned to the “z” axis,
Protein signals: free to precess on the transverse plan
Peak Intensity: Ix Sin(
Delay   max), max=distance of maxima intensity
For example: To observe a peak at 14 ppm at 600 MHz,
=1/[4*(14-4.75)*600.13]=45 usec
Parameter adjustment: Pulprog=p11
p1: 90 pulse, p0: 90 degree “return” pulse, adjust p0 to be slightly shorter than p1 (0.1-0.3 usec).
d19: d19= (1/(2*d)), d = distance of next null (in Hz)
d1 
90y
90-y
Acqu
P. Plateau et and M. Gueron, al., J. Am. Chem. Soc.
1982, 104, 7310-7311

## Slide 21: Binominal: 1-3-3-1
1-3-3-1
a*P1
Acqu
d1
b*P1
b*P1
  
a*P1
a=0.125, b=0.375
P.J. Hore, J. Magn. Reson., 55, 283-300 (1983)
Delay  = 1/(2max)=1/d,
max=distance of maximal intensity d=distance of next null
For example: To observe a peak at 14 ppm at 600 MHz,
=1/[2*(14-4.75)*600.13]=90 usec
Parameter adjustment:
  Pulpro=p1331
  d19: delay for binomial water suppression
  d19 = (1/d), d = distance of next null (in Hz)=2*distance of maximal intensity
  d19= as defined above

## Slide 22: Jump-Return 1-1 and Binominal 1-3-3-1
1-1
1331
1331
1-1
1331
Binomial excitation profiles of 1-1 and 1-3-3-1.
John Cavanagh et al., page 154, “Protein NMR Spectroscopy: Principles and Practice”, Academic Press (1995)

## Slide 23: Jump-Return 1-1 and Binominal 1-3-3-1
Both are for observing fast exchanging protons.
1-3-3-1: Better water suppression (higher receiver gain), but with offset-dependent phase distortion
1-1: low receiver gain, the dispersive tail of water interferes with the signals of interest.
1-3-3-1 (rg=656)
1-1
rg=64

## Slide 24: What are you trying to detect ?
1-3-3-1
WATERGATE
Presaturation

## Slide 25: Water Suppression via Coherence Pathway Rejection
Coherence pathway selected by gradients:
In a gradient selection experiment (echo/antiecho), the water coherence is not “refocused” by the refocus gradient (therefore, is not selected), this naturally suppression the water signal.
Example: cosydfetgp.1, hsqcetf3gp
Figure from John Cavanagh et al.,“Protein NMR Spectroscopy: Principles and Practice”, Academic Press (1995)

## Slide 26: Practical Implementation: 1D, 2D and 3D
Fast exchangeable proton (His sidechain, -OH): 1-1 (good for 2D), 1-3-3-1 (not suitable for 2D, 3D).
Exchangeable NH: Water-flip-back HSQC, Fast-HSQC.
Signals (H) near water: (i.e. TOCSY, COSY) WATERGATE with selective pulse, echo-antiecho PFG.